Purpose: A quantity of 250 pieces of the 724-F-1 boards were slotted to make them easier to install on mainframes.  However, the standoffs that the board mounts to now can fit through the slots so a nylon spacer must be added to the standoffs before installing an F-1 board.  This procedure will only apply when replacing an existing board with a new one.  All units shipped from the factory with the slotted holes have the spacers.

Parts Required:	707-240-099	Qty 11

Description:

- Before installing 724-F-1 board verify if slots on board have been slotted.  If they are round and not slotted this procedure does not apply.

- On each of eleven standoff’s that the F-1 board will mount to, place a drop of hot glue around the top of the standoff and press and nylon spacer over it.

- Mount the F-1 board as usual.

- Program the contents of the EEPROM using EE_EDIT.