Purpose:  To erase the contents of non-volatile memory for security purposes.  Note:  There may be waveforms stored on the hard drive which this procedure will not address.

Parts Required:

None

Warning:  Disassembly described in this section is performed with power removed from the instrument.  Do not remove or replace any of the circuit board assemblies without first verifying that power has been removed from the instrument.  Anti-static protection measures should be taken prior to the handling of any printed circuit boards.

Procedure:

- Disconnect power cable from the instrument.

- Place the LS140 face down into the plastic front panel protective cover that was shipped with the unit or on a soft surface to protect the BNC’s and front panel knob.

- Remove the four screws attaching the rear feet to the rear bezel.

- Remove the six screws attaching the cabinet to the rear bezel.

- Hold the cabinet by the two sides.  Carefully lift the cabinet using a left and right rocking motion.

- Locate the 724-F-5 (Board on left as view from the front.

- Using a small flat blade screwdriver remove U5028 from the socket.  (See diagram below)

- Wait a few seconds and re-install U5028 into the socket being careful not to bend any pins.

- Reassemble cabinet by reversing steps 5 - 2 above.