Description:	This document contains the procedure needed to view and modify the contents of the 			LS140 configuration stored in the EEPROM located on the 724-F-1 board.  Prototype 			units (ie. S/N 50-XXX-XXX) do not have the ability to store a system configuration.

A) To view the current configuration programmed into the EEPROM perform the following steps:

1) Simultaneously press the MENU RETURN, CHANNEL 1, and AUTO SETUP keys 			on the front panel.

2) The test menu should now be displayed.

3) Select Status using the screen menu keys.

4) Select EEprom Dump using screen menu keys.

5) The contents of the EEPROM will be displayed as shown below in FIG 1.

EEprom Contents

MODEL#:   LS140

LS140/U1001

F1 S/N:  [724-F-1/26            ]     [  23 ]     [ B]     [  23 ]     [ 1101 ]

Power  Cycles:  120  	Hours  Up:  503

W/O         REV          S/N             ECO

01	:       [724-F-5                            ]      [    24  ]      [   B  ]      [    25  ]       [   1100  ]

02	:       [200906002                       ]      [      1  ]      [   -   ]      [    15  ]       [   1000  ]

03	:       [724-F-7                            ]      [    24  ]      [   A ]      [     23  ]       [  1002   ]

04	:       [330003202                       ]      [      1  ]      [   -   ]      [    12  ]       [   1000  ]

05	:       [334280067                       ]      [      1  ]      [   -   ]      [    67  ]       [   1000  ]

06	:       [315160110                       ]      [      1  ]      [   -   ]      [    34  ]       [   1000  ]

07	:       [315160200                       ]      [      1  ]      [   -   ]      [    39  ]       [   1000  ]

08	:       [321210099                       ]      [      1  ]      [   -   ]      [    78  ]       [   1000  ]

09	:       [334000386                       ]      [      1  ]      [   -   ]      [    59  ]       [   1000  ]

10	:       < EMPTY >

11	:       < EMPTY >

12	:       < EMPTY >

13	:       < EMPTY >

14	:       < EMPTY >

15	:       < EMPTY >

16	:       < EMPTY >

FIG. 1

LS140 EEprom Dump Screen

B) To edit and view the current configuration programmed into the EEPROM perform the following steps:

1) Insert the EE_EDIT program disk into the floppy drive of the LS140.

2) Turn the unit off and insert a PC/AT type keyboard.

3) Power-up the unit and boot the scope off of the floppy.

4) The EE_EDIT program will automatically load and run.  The screen should appear similar to

FIG. 2 below:

LeCroy  LS140 Manufacturing Support Program    (EE_EDIT  v0.3   DTI '93)

F1:Test EEprom      	F4:Enter F1 S/N    	F7:Add Assy Record

F2:Write Defaults   	F5:Change F1 ECO   	F8:Remove Assy Record

F3:Enter Sys S/N    	F9:Change Assy ECO       F10:Quit

Model            : 		[  LS140                  ]

Sys Serial Number: 	[  LS140/U1001                  ]

W/O   	Rev   	S/N    	ECO

F1 Serial Number : [ 724-F-1           ] 	[  12 ] 	[ B ] 	[  12  ] 	[ 1100 ]

Power Cycles     :	[         23]

Hours Up         : 	[       256]

Assy Record 00   : [ 724-F-5           ] 	[  12  ] 	[ B ] 	[ 123 ] 	[ 1102 ]

Assy Record 01   : [ 334000386      ] 	[  12  ] 	[  -  ] 	[   45 ] 	[ 1000 ]

Assy Record 02   : [ 200356762      ] 	[  10  ] 	[  -  ] 	[   23 ] 	[ 1000 ]

Assy Record 03   : [ 334000386      ] 	[  12  ] 	[  -  ] 	[   45 ] 	[ 1000 ]

Assy Record 04   : [ 350160110      ] 	[  22  ] 	[  -  ] 	[   12 ] 	[ 1000 ]

Assy Record 05   : [ 200906002      ] 	[  13  ] 	[  -  ] 	[   16 ] 	[ 1000 ]

Assy Record 06   : [ 334280067      ] 	[  26  ] 	[  -  ] 	[   15 ] 	[ 1000 ]

Assy Record 07   : [ 200906002      ] 	[  23  ] 	[  -  ] 	[   55 ] 	[ 1000 ]

Assy Record 08   : [                        ]    	[        ] 	[     ] 	[        ] 	[          ]

Assy Record 09   : [                        ]    	[        ] 	[     ] 	[        ] 	[          ]

Assy Record 10   : [                        ]    	[        ] 	[     ] 	[        ] 	[          ]

Assy Record 11   : [                        ]    	[        ] 	[     ] 	[        ] 	[          ]

Assy Record 12   : [                        ]    	[        ] 	[     ] 	[        ] 	[          ]

Assy Record 13   : [                        ]    	[        ] 	[     ] 	[        ] 	[          ]

Assy Record 14   : [                        ]    	[        ] 	[     ] 	[        ] 	[          ]

Assy Record 15   : [                        ]    	[        ] 	[     ] 	[        ] 	[          ]

FIG. 2

EE_EDIT SCREEN DISPLAY

5) If the ECO level needs to be changed press <F5> for 724-F-1 board or <F9> for any other 		    assembly.  The program will prompt for an ECO number.  Enter only the last two digits of 		    the ECO number!

6) If you are replacing an assembly with an assembly with the same part number simply scan the 		    bar code on the new assembly or type in the barcode exactly as it appears on the barcode 		    excluding the "*"s and press <ENTER>.  The program will prompt for an ECO number. Enter 	    only the last two digits of the ECO number!

7) If you are replacing an assembly with one that has a different part number press <F8>.  Enter 		    the assembly record number that contains the part number of the part to be replaced.  Scan the 		    bar code on the new assembly or type in the barcode exactly as it appears on the barcode

excluding the "*"s and press  <ENTER>.  The program will prompt for an ECO number.  		    Enter only the last two digits of the ECO number!