Purpose:  If the hard disk must be replaced on an LS140 with option L1 some setup is required before installing the firmware.

Parts Required:

335830010	qty 1	127 Mb Hard drive

(current approved available drives are Quantum QM303200EXA and Quantum QM31280TM-A)

Software Required:

FDISK

LS140-L1 install software

(LS140-L1-NET is required if unit has network option installed)

Description:

- Connect hard drive to LS140.  Connect keyboard at rear connector of LS140.

- Insert disk label FDISK

- Turn on the power and follow the menu prompts to create a DOS partition (the LS140 may beep several times and reset before booting off of the hard drive, this is required by the routine which programs the CMOS with the correct information about the hard drive).  This partition must be formatted and it must be labeled as LEC_HDD (it is case sensitive).

- Insert disk labeled LS140-L1 install disk and cycle power, follow the on screen instructions.