## 7.6		Calibration Procedures

The following section includes the adjustments required for the power supply, display and front end.  It is recommended that they be verified at one year intervals.

## 7.6.1	93XX-PS1715 Power Supply Calibration

The four voltages are adjustable by ± 5% of the nominal value.

The reference for the measurements are the pins on top of connector J31 located on the 	main board F9384-3.

For the power supply calibration proceed as follow:

§ Turn off the power

§ Remove the top cover (7.2.1)

§ Remove the front frame assembly (7.2.9) and put it to the right of the unit.

§ By using two extension cables, reconnect the processor board to the front panel (J4) 	  	  and to the deflection board (J6).

§ Once the top cover is removed and the front panel is disassembled from the scope, extra 	   cooling of the main board is required. It's mandatory to disconnect the existing Fan 	   	   from connector J33, located on F9384-3 card, and to use a Fan with the air flow 	   	   oriented to the front end section of the board.

§ The front frame assembly is now reconnected to the processor through the extension 	  	   cables.

§ Turn on the power, set the scope to Auto Trigger, and perform the adjustments to get on 	   J31 :

10

9	Pin 3, 5	:	+ 5.10 V ( Min = + 5.05 V, Max. = + 5.2 V )

8	Pin 7, 8	:	- 5.2 V ( Min = - 5.15 V, Max. = - 5.25 V )

7	Pin 9	:	+15 V ( Min = +14.9 V, Max. = +15.1 V )

6	Pin 10	:	- 15 V ( Min = -14.9 V, Max. = -15.1 V )

5	Pin  2, 4, 6:	Ground

4

3

2

1

front

The four potentiometers R108, R206, R306, and R407 are accessible from the front  through holes in the 93XX-PS1715 power supply chassis.

§ Turn the potentiometer clockwise to increase the tension or counterclockwise to 	  	   decrease the voltage. When the adjustment is done, stop the acquisition by depressing the stop trigger push button, and verify that there is no large difference on the + 5.10 V, typically less than 30 mV.

Figure 7.9 : Power Supply

## 7.6.2	93XX-Display Adjustment Procedure

## 7.6.2.1	Introduction

There are a total of 12 potentiometers or variable coils to adjust the deflection and video board.

Video: (2 adjustments)

§ Threshold	:    Level of the video board.

§ Gain		:    Intensity of the screen.

Video board component side

Deflection	:	(10 adjustments)

§ Vosc		:	Frequency of the vertical oscillator.

§ Slope		:	Speed of the horizontal ramp.

§ Focus		:	Focus  of the screen.

§ Cut off		:	Cathode ray tube cut off.

§ Quiescent	:	Standby current of the horizontal defl. amplifier.

§ H Linearity	:	Horizontal linearity.

§ H Size		:	Horizontal size (Max. 165mm).

§ H Offset	:	Horizontal position.

§ V Size		:	Vertical size (Max. 120mm).

§ V Offset	:	Vertical position.

Deflection board component side

## 7.6.2.2  Coarse Adjustment

§ Depress display button.

§ Set W’form + text intensity to 0%.

§ Set grid intensity to 0%

§ Turn fully clockwise the intensity potentiometer on the video board.

§ On the video board connect a digital multimeter on test point : W303

§ Adjust threshold potentiometer to get 2 V ± 0.1 V on W303.

§ Set W'form intensity to 100%.

§ Set grid intensity to 60%.

§ Adjust H-size, H-offset, V-size, V-offset to center the image in the screen.

The vertical position should be adjusted to get the push buttons of the front panel in 	  	  front of the software menus, use the utilities set up.

The small magnets mounted on the deflection yoke influence the vertical position.

§ Turn the quiescent potentiometer clockwise until the default of the horizontal lines just 	   disappears from the vertical center of the screen.

§ Increase the cut off until a vertical line appears on the right side of the screen.

§ Adjust the slope potentiometer to get 5 mm gap between the highlighted vertical line 	  	  and the right border of the selection menus.

§ Adjust H-linearity to get the best linearity.

## 7.6.2.3  Fine Adjustment

The final adjustment of the intensity, cut off, and focus must be made in a dark room.

§ Set W'form intensity to 30%.

§ Set grid intensity to 0%.

§ Adjust the cut off potentiometer until the highlighted vertical line disappears from the 	right side 	    of the screen.

§ Set W'form intensity to 20%.

§ Display four traces.

§ On the video board adjust the gain potentiometer (intensity) in order to get the text

just readable.

§ Set W'form + text intensity to 70%.

§ Set grid intensity to 30%

§ Adjust the focus (usually fully clockwise) for most uniform focus

over the entire screen.

§ In a standard luminosity environment set W'form + text to 90%, and

grid intensity to 60%.

§ Verify the intensity, focus, and contrast adjustment, for best

definition of the displayed text.

### C A U T I O N

Never change the Vosc calibration.

## 7.6.3	Front End Test and Calibration Procedure

## 7.6.3.1	Introduction

The adjustments describe in the following calibration procedure require extension of the 	front panel assembly out of the scope, using two flat cables.

In order to access the front end potentiometers and variable caps located underneath the 	Cathode Ray tube and deflection board, dismount the front panel assembly from the scope and reconnect it to the processor board connectors J4 and J6,  using the extension cable set.

Once the top cover is removed and the front panel is disassembled from the scope, extra 	cooling of the main board is required. It's mandatory to disconnect the existing Fan from connector J33, located on F9384-3, and to use a Fan with the air flow oriented to the front end section of the board.

## 7.6.3.2	Power Supplies

Remove the upper shield to access the test points and potentiometers.

Use a high precision DMM.

## 7.6.3.2.1	HWD418 Temperature Verification

§ Put a jumper between pin 1 and 2 of TP19

§ When the HWD418 of Channel 1 is cold, check on TP10 pin 9 that :

Vtemp = min 3.4 Volt, max + 3.7 Volt

§ When the HWD418 of Channel 2 is cold, check on TP14 pin 9 that :

Vtemp = min 3.4 Volt, max + 3.7 Volt

If Vtemp is below 3.0 V the HWD418 must be replaced

§ Remove the jumper from TP19

7.6.3.2.2	HFE620 Power Supply Adjustment

§ Adjust R431 to get on TP8  pin 3 (+) : +14 Volt ( min 13.7 V, max. 14.3 V )

7.6.3.2.3	MWD418 Power Supply Adjustment

§ Adjust R844 to get on TP8  pin 9 (+) : +2.49 Volt ( min 2.47 V, max. 2.51 V )

7.6.3.2.4	Front End Power Supply Verification

§ Check on test point TP6 the following tensions :

§ Pin 1 and Pin 2 : Ground

§ + 12 VFEP : pin 3 (+)  =  + 12 Volt ± 2% : if problem troubleshoot regulator VR2

Figure 7.10 : F9384-3 Main Board Calibration Layout

§ + 12 VFE 	:	pin 5 (+)  =  + 12 Volt ± 2%	  if problem troubleshoot regulator VR4

§ Check on test point TP7 the following tensions :

§ Pin 1 and Pin 2 : Ground

§ + 2 VFE	:	pin 3 (+)  =  + 2 Volt ± 2% 	  if problem troubleshoot regulator VR5

§ + 7 VFEA	:  	pin 4 (+)  =  + 7 Volt ± 2%	  if problem troubleshoot regulator VR8

§ + 5 VFE   	: 	pin 5 (+)  =  + 5 Volt ± 2% 	  if problem troubleshoot regulator VR12

§ + 7 VFEB	:	pin 6 (+)  =  + 7 Volt ± 2%	  if problem troubleshoot regulator VR9

§ + 10 VFE	:	pin 8 (+)  =  + 10 Volt ± 2%	  if problem troubleshoot regulator VR7

§ - 10 VFE	:	pin 7 (-)  =  - 10 Volt ± 2%	  if problem troubleshoot regulator VR6

§ - 6 VFE	:	pin 9 (-)  =  - 6 Volt ± 2% 	  if problem troubleshoot regulator VR11

§ - 5 VFE	:	pin 10 (-)  =  - 5 Volt ± 2% 	  if problem troubleshoot regulator VR10

## 7.6.3.2.4	ADC Power Supply Verification

§ Check on test point TP8 the following tensions :

§ Pin 1 and Pin 2 : Ground

§ + 5 V	:	pin 4 (+)  =  + 5 Volt ± 2% 	  if problem troubleshoot regulator VR14

§ + 12 V	:	pin 5 (+)  =  + 12 Volt ± 2%	  if problem troubleshoot regulator VR15

§ + 8 V	:	pin 6 (+)  =  + 8 Volt ± 2%	  if problem troubleshoot regulator VR13

## § - 12 V	:	pin 7 (-)  =  - 12 Volt ± 2%	  if problem troubleshoot regulator VR16

§ - 4.8 V 	: 	pin 10 (-)  = - 4.8 Volt ± 2%	  if problem troubleshoot U69 & Q38

§ Check on test point TP18 the following tensions :

§ Pin 1 and Pin 2 : Ground

§ + 1.2 V	:	pin 3 (+)  =  + 1.2 Volt ± 2%  if problem troubleshoot U122 & Q76

§ + 2.4 V	:	pin 4 (+)  =  + 2.4 Volt ± 2%  if problem troubleshoot U122 & Q57

§ + 3.6 V 	: 	pin 5 (+)  =  + 3.6 Volt ± 2%  if problem troubleshoot U122 & Q58

§ + 4.8 V 	: 	pin 6 (+)  =  + 4.8 Volt ± 2%  if problem troubleshoot U122 & Q75

§ + 8.4 V 	: 	pin 7 (+)  =  + 8.4 Volt ± 2%  if problem troubleshoot U134 & Q81

§ + 6 V 	: 	pin 8 (+)  =  + 6 Volt ± 2% 	  if problem troubleshoot U134 & Q82

§ + 7.6 V 	: 	pin 9 (+)  =  + 7.6 Volt ± 2%  if problem troubleshoot U134 & Q79

§ + 9 V 	: 	pin 10 (+) =  + 9 Volt ± 2% 	  if problem troubleshoot U103 & Q49

§ Adjust R406 to get on TP10  pin 10 (+) : - 3.32 Volt  ± 2 %

§ Adjust R505 to get on TP14  pin 10 (+) : - 3.32 Volt  ± 2 %

§ Perform a general instrument reset. Simultaneously press the autosetup button, the top 	   menu button and the return button.

§ The scope must be  stopped ( no acquisition ).

§ Check  on pin 3, 4, 15 and 16 of U114 that the voltage is between  minimum .9 V, 	  	   and maximum 1.2 V.

§ Adjust R735 to get on TP15  pin 8 (+) :- 1.84 Volt ( min - 1.83 V, max. : - 1.85 V )

§ Adjust R744 to get on TP16  pin 8 (+) :- 1.84 Volt ( min - 1.83 V, max. : - 1.85 V )

7.6.3.3	16  bit DAC  Calibration

The gain and offset adjustments of the 16 bit DAC should only be performed if the

instrument did not meet the specifications as checked in this procedure.

These adjustments were preset at the factory to their optimum setting.

§ Enter in the internal calibration diagnostics by simultaneously depressing the third

and fourth push buttons, and then by depressing the fifth.

§ Select Development menu

§ Select Front End Control menu

§ Set Cal Level  : /1 ( 0.00 V )

§ Check that we get on pin 7 ( + ) of TP3 : 0.00 mV  ± 0.2 mV

If it is determined that the offset needs to be calibrated :  adjust R875

§ Set Cal Level  /1 to + 5.00 V ( use front panel potentiometer )

§ Check that we get on pin 7 ( + ) of TP3 : - 500 mV ±  2%

§ If it is determined that the gain needs to be calibrated : adjust R874

§ Set Cal Level  /1 to + 8.00 V ( use front panel potentiometer )

§ Check that we get on pin 7 ( + ) of TP3 : - 800 mV ±  2%

§ Set Cal Level  /1 to - 5.00 V ( use front panel potentiometer )

§ Check that we get on pin 7 ( + ) of TP3 : + 500 mV ±  2%

§ Set Cal Level  /1 to - 8.00 V ( use front panel potentiometer )

§ Check that we get on pin 7 ( + ) of TP3 : + 800 mV ±  2%

§ Reset Cal Level  to : --

## 7.6.3.4	Time Base Calibration

These adjustments were preset at the factory to their optimum setting and further 	adjustment may result in reduced instrument performance.

7.6.3.4.1	100 MHz Clock Calibration

Setup a levelled sine wave generator and connect the output to Channel 1

§	Frequency	:	100.02  MHz

§	Turn on trace	:	Ch1

§	Input gain	:	.1 V/div

§	Trigger setup	:	Edge

§	Trigger on	:	1

§	Coupling 1	:	DC

§	Slope 1	:	Pos

§	Mode	:	Auto

§	Holdoff	:	Off

§	Delay	:	10 %

§	Timebase	:	20 msec/div

§	Sampling	:	Single Shot ( 100 MS/s )

§	Record up to	:	25 K

§	Adjust the generator output amplitude and Ch1 offset to get a six divisions peak to  			peak amplitude sine wave.

§	Press 	:	Cursors/Measure

§	Measure	:	Parameters

§	Mode	:	Custom

§	Statistics	:	On

§	Change parameters

§	On line 1

§	Category	:	All

§	Measure 	:	Freq of 1

§ Check that : average Freq(1) readout is 20 KHz ± 2.5 % (min 19.5 KHz, max 20.5 KHz)

§ If it is determined that the 100 MHz clock needs to be calibrated : adjust C487

7.6.3.4.2	833 MHz Clock Calibration

Setup a levelled sine wave generator and connect the output to Channel 1

§	Frequency	:	250.9  MHz

§	Turn on trace	:	Ch1

§	Input gain	:	.1 V/div

§	Trigger setup	:	Edge

§	Trigger on	:	1

§	Coupling 1	:	DC

§	Slope 1	:	Pos

§	Mode	:	Auto

§	Delay	:	10 %

§	Timebase	:	.2 msec/div

§	Sampling	:	Single Shot ( 250 MS/s )

§	Record up to	:	25 K

§	Adjust the generator output amplitude and Ch1 offset to get a six divisions peak to  			peak amplitude sine wave.

§	Press 	:	Cursors/Measure

§	Measure	:	Parameters

§	Mode	:	Custom

§	Statistics	:	On

§	Change parameters

§	On line 1

§	Category	:	All

§	Measure 	:	Freq of 1

§ Check that : average Freq(1)  readout is 1 MHz ± 5 % (min 950 KHz, max. 1050 KHz)

§ If it is determined that the 833 MHz clock needs to be calibrated : adjust C245

7.6.3.5	HFE 620 Calibration

7.6.3.5.1	Channel 1 HFE620 Offset Calibration

With no signal connected to the inputs set 9384 DSO settings as follow :

§	Input gain	:	50 mV/div

§	Timebase	:	.2 msec/div

§	Trigger mode	:	Auto

§	Sampling	:	Single Shot ( 250 MS/s )

§ Enter in the internal calibration diagnostics by simultaneously depressing the third

and fourth push buttons, and then by depressing the fifth.

§ Select Development menu

§ Calibration control

§ Recalibration 	:	OFF

§ Set 0.0 mV  ± 500 mV on TP4 pin 9  ( pin 42 of U1 Hybrid HFE620 ) with front panel 	   Channel 1 offset potentiometer.

§ Adjust potentiometer R5 to get 0.0 mV ± 2 mV on TP4 pin 7

§ Check on TP4 pin 9 : 0.0 mV ± 500 mV

§ Adjust potentiometer R9 to get 0.0 mV ± 2 mV on TP9 pin 3 (pin H2 of U81 HWD418)

§ Check on TP4 pin 9 : 0.0 mV ± 500 mV and on TP9 pin3 : 0.0 mV ± 2 mV ( trim R9 )

§	Timebase	:	.5 msec/div

§	Trigger mode	:	Auto

§	Sampling	:	Single Shot ( 100 MS/s )

§ Check on TP4 pin 9 : 0.0 mV ± 500 mV

§ Adjust potentiometer R13 to get 0.0 mV ± 2 mV on TP9 pin 3 ( HWD418 )

§ Check on TP4 pin 9 : 0.0 mV ± 500 mV

§ Adjust potentiometer R15 to precisely align the trace of Channel 1 to the center vertical

graticule line.

7.6.3.5.2	Channel 2 HFE620 Offset Calibration

With no signal connected to the inputs set 9384 DSO settings as follow :

§	Input gain	:	50 mV/div

§	Timebase	:	.2 msec/div

§	Trigger mode	:	Auto

§	Sampling	:	Single Shot ( 250 MS/s )

§ Enter in the internal calibration diagnostics by simultaneously depressing the third

and fourth push buttons, and then by depressing the fifth.

§ Select Development menu

§ Calibration control

§ Recalibration 	:	OFF

§ Set 0.0 mV  ± 500 mV on TP4 pin 10  ( pin 42 of U2 Hybrid HFE620 ) with front panel 	   Channel 2 offset potentiometer.

§ Adjust potentiometer R6 to get 0.0 mV ± 2 mV on TP4 pin 8

§ Check on TP4 pin 10 : 0.0 mV ± 500 mV

§ Adjust potentiometer R10 to get 0.0 mV ± 2 mV on TP11 pin 3 ( pin H2 of U82 )

§ Check on TP4 pin 10 : 0.0 mV ± 500 mV and on TP11 pin3 : 0.0 mV ± 2 mV (trim R10)

§	Timebase	:	.5 msec/div

§	Trigger mode	:	Auto

§	Sampling	:	Single Shot ( 100 MS/s )

§ Check on TP4 pin 10 : 0.0 mV ± 500 mV

§ Adjust potentiometer R14 to get 0.0 mV ± 2 mV on TP11 pin 3 ( HWD418 )

§ Check on TP4 pin 10 : 0.0 mV ± 500 mV

§ Adjust potentiometer R16 to precisely align the trace of Channel 2 to the center vertical

graticule line.

§ Select Development menu

§ Calibration control

§ Recalibration 	:	On

§ Press Reset Calibration Curves

7.6.3.5.3	Channel 1 HFE20 Trigger Hysteresis Adjustment

Setup a levelled sine wave generator.

§	Frequency	:	100  KHz

§	Connect the generator output to Channel 1

§	Turn on trace	:	Ch1

§	Display setup	:	Standard, Persistence off, Dot join on, Single grid

§	Input gain	:	50 mV/div

§	Input offset	:	0.0 mV

§	Trigger setup	:	Edge

§	Trigger on	:	1

§	Coupling 1	:	DC

§	Slope 1	:	Pos

§	Mode	:	Normal

§	Holdoff	:	Off

§	Delay	:	10 %

§	Timebase	:	20 msec/div

§	Sampling	:	Single Shot ( 100 MS/s )

§	Channel use 	:	2

§	Record up to	:	25 K

§	Adjust the generator output amplitude to get one division peak to  					peak amplitude sine wave ( -22 dBm ).

§	Set Trigger Level 	: 0 mV

§	Check that the 9384 DSO triggers at DC 0 mV

§	Set Trigger Level 	: -2 mV

§	Adjust potentiometer R7 such that the DSO doesn’t trigger at  DC -2 mV ( slow 			trigger and no flashes in box next to normal ) but triggers at DC 0 mV.

§ Check that the scope triggers from DC 0mV to +25mV, but doesn’t trigger below -1mV

7.6.3.5.4	Channel 2 HFE20 Trigger Hysteresis Adjustment

Setup a levelled sine wave generator and connect the output to Channel 2

§	Frequency	:	100  KHz

§	Turn on trace	:	Ch2

§	Input gain	:	50 mV/div

§	Input offset	:	0.0 mV

§	Trigger setup	:	Edge

§	Trigger on	:	2

§	Coupling 2	:	DC

§	Slope 2	:	Pos

§	Mode	:	Normal

§	Delay	:	10 %

§	Timebase	:	20 msec/div

§	Sampling	:	Single Shot ( 100 MS/s )

§	Record up to	:	25 K

§	Adjust the generator output amplitude to get one division peak to  					peak amplitude sine wave ( -22 dBm ).

§	Set Trigger Level 	: 0 mV

§	Check that the 9384 DSO triggers at DC 0 mV

§	Set Trigger Level 	: -2 mV

§	Adjust potentiometer R8 such that the DSO doesn’t trigger at  DC -2 mV ( slow 			trigger and no flashes in box next to normal ) but triggers at DC 0 mV

§ Check that the scope triggers from DC 0mV to +25mV, but doesn’t trigger below -1mV

7.6.3.5.5	Channel 1 HFE620 Flatness Calibration

§	Connect the Probe Calibrator output to Channel 1, using a 5 nsec BNC cable

§	Select	:	Utilities

§	Press	:	Cal BNC Setup

§	Mode	:	Cal signal

§	Shape	:	Square

§	Amplitude	:	.5 V

§	Set Frequency	:	500 Hz

§	Turn on trace 	:	Ch1

§	Display setup	:	Standard, Persistence off, Dot join on, Single grid

§	V/div. offset	:	Normal

§	Input gain 	: 	50 mV/div.

§	Input offset	:	- 130 mV

§	Trigger setup	:	Edge

§	Trigger on	:	1

§	Coupling 1	:	DC

§	Slope 1	:	Neg

§	Trigger level	:	DC 130 mV

§	Mode	:	Normal

§	Delay	:	10 % Pre-Trigger

§	Timebase	:	.1 msec/div.

§	Channel use	:	2

§	Record up to	:	25K

§	Turn on trace	:	A

§	Select Math Setup

§	For Math	:	Use at most 5000 points

§	Redefine A

§	Use Math ?	:	Yes

§	Math Type 	:	Average

§	Avg Type	:	Continuous with 1:7 weighting or Summed

§	Of	:	Channel 1

§	With the vertical zoom set A to 5 mV

§	Use the vertical zoom position knob to precisely align trace A to the center of screen

at nine horizontal divisions after the trigger.

§	Timebase	:	1 msec/div

§	Sampling	:	RIS ( 2.5 GS/s )

§	Cursors/Measure	:	Cursors

§	Mode	:	Time

§	Type	:	Relative

§	Use the reference cursor position knob, to move the reference marker at the trigger  			point.

§	Use the difference cursor position knob, to move the difference marker at 500 ns after 		the trigger point.

§	Adjust R1 so that trace A is at -1 vertical division from the center of screen at the 			difference marker position.

§	During the adjustment push reset calibration curves, from the Internal Calibration 			Diagnostics, Development, Calibration Control menus.

§	Timebase	:	.1 msec/div.

§	Check  that trace A is precisely aligned to the center of screen at nine horizontal 			divisions after the trigger point ( use vertical zoom position knob ).

§	Timebase	:	1 msec/div

§	With the horizontal zoom set A to 20 ns

§	Use the horizontal zoom position knob to expand the trailing edge of  trace 1

§	At the difference marker position, check that trace A  is between 0 and  +1 vertical 			division from the center of screen

7.6.3.5.6	Channel 2 HFE620 Flatness Calibration

§	Connect the Probe Calibrator output to Channel 2, using a 5 nsec BNC cable

§	Select	:	Utilities

§	Press	:	Cal BNC Setup

§	Mode	:	Cal signal

§	Shape	:	Square

§	Amplitude	:	.5 V

§	Set Frequency	:	500 Hz

§	Turn on trace 	:	Ch2

§	V/div. offset	:	Normal

§	Input gain 	: 	50 mV/div.

§	Input offset	:	- 130 mV

§	Trigger setup	:	Edge

§	Trigger on	:	2

§	Coupling 2	:	DC

§	Slope 2	:	Neg

§	Trigger level	:	DC 130 mV

§	Mode	:	Normal

§	Delay	:	10 % Pre-Trigger

§	Timebase	:	.1 msec/div.

§	Record up to	:	25K

§	Turn on trace	:	B

§	Select Math Setup

§	For Math	:	Use at most 5000 points

§	Redefine B

§	Use Math ?	:	Yes

§	Math Type 	:	Average

§	Avg Type	:	Continuous with 1:7 weighting or Summed

§	Of	:	Channel 2

§	With the vertical zoom set B to 5 mV

§	Use the vertical zoom position knob to precisely align trace B to the center of screen

at nine horizontal divisions after the trigger.				 .

§	Timebase	:	1 msec/div

§	Sampling	:	RIS ( 2.5 GS/s )

§	Cursors/Measure	:	Cursors

§	Mode	:	Time

§	Type	:	Relative

§	Use the reference cursor position knob, to move the reference marker at the trigger  			point.

§	Use the difference cursor position knob, to move the difference marker at 500 ns after 		the trigger point.

§	Adjust R2 so that trace B is at -1 vertical division from the center of screen at the 			difference marker position.

§	During the adjustment push reset calibration curves, from the Internal Calibration 			Diagnostics, Development, Calibration Control menus

§	Timebase	:	.1 msec/div.

§	Check  that trace B is precisely aligned to the center of screen at nine horizontal 			divisions after the trigger point ( use vertical zoom position knob ).

§	Timebase	:	1 msec/div

§	With the horizontal zoom set B to 20 ns

§	Use the horizontal zoom position knob to expand the trailing edge of  trace 2

§	At the difference marker position, check that trace B  is between 0  and  +1 vertical 			division from the center of screen.

7.6.3.6	HWD418 Calibration

7.6.3.6.1	Channel 1 HWD418 Calibration

With no signal connected to the inputs set 9384 DSO settings as follow :

§	Turn on trace	:	Ch1

§	Input gain	:	50 mV/div

§	Input offset	:	0 mV

§	Timebase	:	.1 msec/div

§	Trigger mode	:	Auto

§	Sampling	:	Single Shot ( 500 MS/s )

§	Display setup	:	Standard, Persistence off, Dot join on, Dual grids

§	Turn on trace	:	A

§	Vertical offset of A	:	0 mV

§	Select Math Setup

§	Redefine A

§	Use Math ?	:	No

§	Trace A is Zoom of:	Channel 1

§	With the horizontal zoom knob set A to 5 nS

§	Enter in the internal calibration diagnostics by simultaneously depressing the third

and fourth push buttons, and then by depressing the fifth.

§	Select Development menu

§	Select MWD Digitizer Control

§	Correction Nsf

§	Select MWD Special Modes

§	Control Sampling

§	R1C Sync On

§	1st Sample 10

§	No. Samples 4

§	Control Cal/Avg

§	Dyn Cal State On

§	DC Avg 8

§	Verify Avg. 8

§	Control Cell

§	Row 2

§	Column 6

§	Control Tricks

§	Tricks Q/Vtest

§ Select Front End Control menu

§ Set Cal Level  : /1

§ Adjust Cal Level to get 80.00 mV ± 2 mV on TP9 pin 3 ( pin H2 of  HWD418 )

§	Adjust R597 and C385 to get the last four dots of the plateau as 					close as possible of a straight flat line, and below + 2 vertical divisions.

§	Adjust Cal Level to get -80.00 mV ± 2 mV on TP9 pin 3 ( pin H2 of  HWD418 )

§	Check that the last four dots of the plateau are close of a straight flat 					line, and above - 2.5 vertical divisions.

7.6.3.6.2	Channel 2 HWD418 Calibration

With no signal connected to the inputs set 9384 DSO settings as follow :

§	Turn on trace	:	Ch2

§	Input gain	:	50 mV/div

§	Input offset	:	0 mV

§	Timebase	:	.1 msec/div

§	Trigger mode	:	Auto

§	Sampling	:	Single Shot ( 500 MS/s )

§	Display setup	:	Standard, Persistence off, Dot join on, Dual grids

§	Turn on trace	:	B

§	Vertical offset of B	:	0 mV

§	Select Math Setup

§	Redefine B

§	Use Math ?	:	No

§	Trace B is Zoom of:	Channel 2

§	With the horizontal zoom knob set B to 5 nS

§	Enter in the internal calibration diagnostics by simultaneously depressing the third

and fourth push buttons, and then by depressing the fifth.

§	Select Development menu

§	Select MWD Digitizer Control

§	Correction Nsf

§	Select MWD Special Modes

§	Control Sampling

§	R1C Sync On

§	1st Sample 10

§	No. Samples 4

§	Control Cal/Avg

§	Dyn Cal State On

§	DC Avg 8

§	Verify Avg. 8

§	Control Cell

§	Row 2

§	Column 6

§	Control Tricks

§	Tricks Q/Vtest

§ Select Front End Control menu

§ Set Cal Level  : /1

§ Adjust Cal Level to get 80.00 mV ± 2 mV on TP11 pin 3 ( pin H2 of  HWD418 )

§	Adjust R610 and C403 to get the last four dots of the plateau as close as possible

of a straight flat line, and below + 2 vertical divisions.

§	Adjust Cal Level to get -80.00 mV ± 2 mV on TP11 pin 3 ( pin H2 of  HWD418 )

§	Check that the last four dots of the plateau are close of a straight flat 					line, and above - 2.5 vertical divisions.