# SECTION	6	MAINTENANCE

## 6.1	Introduction

This section contains information necessary to disassemble, assemble, maintain, calibrate 	and troubleshoot the LeCroy 9384 digital storage oscilloscope.

## 6.2	Disassembly and Assembly Procedure

The disassembly and assembly procedures detailed below refer to the assembly and 	disassembly diagram 6.2.3, and the view of figures 6.1, 6.2, 6.3, 6.4, 6.5, 6.6 and 6.7.

Please study the diagram and figures before attempting disassembly.

### W A R N I N G

Before removing any parts from the LeCroy 9384, be sure to read carefully the 	instructions referring to those parts, noting any precautions needed to avoid problems 	caused by mechanical behavior, high voltage supplies, etc.

### C A U T I O N

The usual precautions against static electricity are required, (see 1.10)

## 6.2.1	Removal of the Upper Cover (5.10)

The top cover (5.9) is secured by two M4x5 screws (5.11) on both sides of the front 	panel assembly (2), and by two M4x8 screws (5.10) on the rear panel (3). Remove the 	screws and carefully slide the cover off the unit to the rear. Removal of the top cover 	gives access to the boards and parts listed in section 6.2.3.

## 6.2.2	Removal of the PS9384 Power Supply (4)

### W A R N I N G

Ensure the line cord is disconnected. Remove the following:

§	Top cover (6.2.1).

§	One M4X8 screw (5.2) from left side of the bottom cover (1.1).

§	Two M4X8 screws (5.1) from left side of the rear panel (3).

Disconnect the following:

§	Base card power cable (5.9) from 9384-3 main board connector J31.

§	PS9384 line input cable (3.20) from connector J6.

The power supply can now be removed vertically from the oscilloscope.

## 6.2.3	Disassembly and Assembly Diagram

Disassembly : If it becomes necessary to replace a board or a part, use the 	disassembly diagram to disassemble the unit. Any board can be removed if items 	higher in the diagram and connected by a line are already out.

Assembly : Reassemble the unit in the reverse order.

## Figure 6.1 : 9384 Assembly

## Figure 6.2 : Lower Cover Assembly

Figure 6.3 : Front Frame Assembly

## Figure 6.4 : Rear Panel Assembly

## 6.2.4	Removal of the F9300-4 GPIB/RS232 Interface (3.10)

The GPIB/RS232 interface (3.7) is vertically mounted on the rear panel (3.1).

Remove the following:

§	Top cover (6.2.1).

§	Two M3x6 screws (3.17) and washers from the rear panel (3.1).

§	Disconnect the flat cable (3.14) from the processor board (1.25) connector J5.

The GPIB/RS232 board can be removed forwards from the rear panel.

6.2.5	Removal of the Fan (3.3)

Remove the following:

§	Top cover of 9384 (6.2.1)

§	Four screws (3.4)  from the rear panel (3.1).

§	Disconnect the fan power cable from the main card F9384-3 connector J33.

The fan (3.3) part number : 7093XX902 can be removed from the unit.

### C A U T I O N

Note the air flow, the fan extracts air from the unit and expels it.

6.2.6	Removal of the Fuse Holder (3.6)

### W A R N I N G

Disconnect the power cord.

Remove the following:

§	Top cover (6.2.1).

§	Two screws (3.6) from the rear panel.

§	Disconnect the power cable from the power supply.

§	Disconnect the earth cable.

The fuse holder assembly (3.5) can be removed from the rear panel (3.1).

6.2.7	Removal of the 93XX-Video (2.20)

§	Remove the top cover (6.2.1).

§	Disconnect the ground cable from CRT (black wire)

§	Disconnect the monitor cable (2.21) from the deflection board, connector

W301 & W302.

Ease the video board (2.20) carefully toward the back of the DSO, until it is free.

6.2.8	Removal of the 93XX-Yoke (2.19)

§ Remove the top cover (6.2.1).

§ Remove the 93XX-video board (6.2.7)

§ Disconnect the cable from the deflection board connector W201.

§ Loose the screw on the yoke ring holder.

The deflection yoke (2.19) can be removed from the cathode ray tube (2.14).

6.2.9	Removal of the front frame Assembly (2)

Remove the following:

§ Top cover (6.2.1)

§ Two screws (5.8) that secure the front frame assembly (2) to the lower cover (1.1).

§ Disconnect the front panel flat cable (2.11) from the processor (1.25) connector J4.

§ Disconnect the deflection flat cable (2.22) from the processor board connector J6.

The front frame assembly (2) with the CRT (2.14), yoke (2.19), video (2.20), deflection 	(2.17), front panel (2.7) and keyboard (2.6) can with care be removed forward from the 	unit.

### C A U T I O N

Hold the CRT very carefully, or place soft padding under it.

6.2.10	Removal of the 93XX-Deflection (2.17)

The deflection board (2.17) is situated to the back of the front panel (2.5).

Remove the following

§ Top cover (6.2.1).

§ Front frame assembly (6.2.9).

§ Disconnect the monitor cable (2.21) which leads to the video board (2.20), connector 	  	   W301 and W302.

§ Disconnect the cable from the deflection yoke, connector W201.

§ Disconnect the red high voltage cable from the receptacle at the right side of the CRT 		   (2.14).

### W A R N I N G

Touch the free end of the high voltage cable to the ground, this ensures that no significant 	charge remains. The CRT must be discharged similarly, using a tool or a long screw 	driver which is first placed to the ground and on the CRT receptacle.

Remove the four M35x10 screws (2.18) that secure the deflection board to the plastic 	front frame.

The board (2.17) can now be removed from the unit.

6.2.11	Removal of the 93XX-CRT (2.14)

It is necessary to remove the front frame assembly (6.2.9). The CRT is secured to the 	plastic front frame by four screws (2.16).

§ Remove the 93XX-video (6.2.7).

§ Remove the 93XX-yoke (6.2.8).

§ Disconnect the red high voltage cable from the deflection board.

§ Remove the four screws.

The CRT can now be removed from the front frame.

### W A R N I N G

Use care when handling the CRT. Avoid striking it on any object which may cause the 	tube to implode. Store the cathode ray tube face down on a soft surface.

To avoid electrical shock the CRT should be discharged after the 9384 oscilloscope is 	powered OFF. After disconnecting the red high voltage cable, ground the cable to the 	metallic display support, repeat the operation to fully dissipate the charge.

6.2.12	Removal of the F9354-5 Front Panel (2.5)

Remove the following:

§ Upper cover (6.2.1).

§ Front frame assembly (6.2.9).

§ 93XX-deflection board (6.2.10).

§ Four screws (2.12) that secure the front panel.

The front panel (2.5) with the keyboard (2.6) can be removed forward from the unit.

6.2.13	Removal of the Front Panel Keyboard (2.6)

Remove the following:

§ Upper cover (6.2.1).

§ Front frame assembly (6.2.9).

§ 93XX-deflection board (6.2.10).

§ F9350-5 front panel (6.2.12).

§ The 13 rotary knobs (2.9 and 2.10).Take great care of the soft plastic

§ One screw (2.8) that secures the keyboard to the front panel.

§ Disconnect the flat ribbon cable from the front panel connector J2, and remove the 	  	   keyboard P/N : 729350513.

### C A U T I O N

When removing or installing the keyboard or the front panel, be careful of the 	fragile flat ribbon cable and connector.

6.2.14	Removal of the Processor (1.21)

The processor board F9302-1-8 is located along the right side of the instrument.

Remove the following:

§ Top cover (6.2.1).

§ Front frame assembly (6.2.9).

§ Disconnect the flat cable (3.14) from the F9300-4 GPIB interface connector J5

The processor can be removed vertically from the main card (1.20) connector J34

### C A U T I O N

Static electricity can damage components (RAM, EPROM’s, microprocessor...).

Antistatic precautions are required.

6.2.15	Removal of the F9384-3 Main Card (1.20)

Remove the following:

§ Top cover (6.2.1).

§ Front frame assembly (6.2.9).

§ Power supply (6.2.2).

§ Processor (6.2.14).

The main board with the upper shield (1.2) is horizontally mounted to the lower case 	cover (1.1).

§ Remove the twelve M3x20 screws (5.3) and four M2.5x6 (5.6) that secure the upper 	  	   shield (1.2) to the main board and front panel.

§ Remove the two M4x8 (5.1) and one M3x6 (5.2) that secure the rear panel assembly (3)  	   to the lower cover (1.1)

§ Disconnect the fan cable from connector J33

The upper shield (1.2) attached to the rear panel (3) can be removed  forward from the 	board.

§ Remove the five M3x6 screws (1.22), four M3x16 (1.18) and three M3x6 flat head 	 	   screws (1.4) that secure the board to the lower cover (1.1).

The main board F9384-3 (1.10) with base shield (1.15) and card panel (1.11) can be 	removed from the scope.

### C A U T I O N

Antistatic precautions are required.

6.2.16	Removal of the Handle (1.4)

The handle with two black end caps is secured to the right side of the lower cover (1.1) 	by two screws (1.10) and washers (1.9).

§ Remove the upper cover (6.2.1), and processor board (6.2.14).

The handle can be removed from the lower case.

6.2.17	Removal of the Foot Support (1.8)

The two foot supports are clipped on the lower cover (1.1).

§ Remove the foot (1.5) or the support (1.6) by inserting a small flat screwdriver under 	 	   the support

6.2.18	Removal of the 93XX-FD01 Floppy Disk Drive Option

§ Remove the upper cover ( 6.2.1 ).

§ Disconnect the flat ribbon cable from the F9300-6 interface ( see figure 6.6).

§ Remove the two M3x6 screws that secure the floppy drive support to the upper cover.

§ Remove the support 70FD01021 and frame 70FD01031 from the cover.

§ Remove the four M2.5x4 screws (6.14) that secure the floppy to the support

The floppy disk (6.4) drive can be removed from the frame

6.2.19	Removal of the 93XX-GP01 Graphic Printer and F9300-7 Controller Option

§ Remove the upper cover ( 6.2.1 ).

§ Disconnect the power cable from the auxiliary power supply ( 315070025 )

( see figure 6.3 & 6.7 ).

§ Disconnect the flat ribbon cable ( 780791604 ) from the F9300-7 controller

( see figure 6.7 ).

§ Disconnect the flat ribbon cable ( 780721022 ) between the F9300-6 interface and 	  	   F9300-7 controller.

§ Remove the four M3x6 screws (6.21) that secure the F9300-7 controller to frame

( 70GP01031 ).

§ Remove the F9300-7 controller (6.9).

§ Remove the two M3x6 screws (6.18) that secure the printer to the frame

The graphic printer ( 6.3) can now be removed from the upper cover.

6.2.20	Removal of the F9300-6 Centronics Interface Option

§ Remove the upper cover ( 6.2.1).

§ Remove the two M3x6 screws from the rear panel

§ Disconnect the flat cable from the F9300-4 GPIB/RS232 board ( see figure 6.6 ).

The Centronics interface board can be removed forward from the rear panel.

Figure 6.5 : Floppy Drive Assembly P/N: 335023203

## Figure 6.6 : Graphic Printer Assembly

## 6.3	Software Upgrade Procedure

F9302-1-8 processor board has one 8 MB Flash Prom which contains the program 	memory and the character font used by the graphic processor of the raster scan display.

After any software change, a general instrument reset is mandatory. Simultaneously press 	the autosetup button, the top menu button and the return button.

## 6.3.1	Upgrading Firmware

LeCroy Corporation has a policy of continually improving and upgrading its products.

The 9384 instrument is equipped with Flash Prom on processor board, the Software is 	upgraded to the latest version using either the Memory Card interface or the Floppy disk

drive.

6.3.1.1	Upgrading Firmware from Memory Card

A single memory card  containing the 93XX36XX.bin file in the Lecroy_P directory is 	required.

§ Insert the card and cycle power to the scope. Once re-booted enter the Utilities, Special 	   Modes, Firmware Update menu.

§ Change the Update From control to Card

§ Press Update Program.

The Software is then downloaded to the Flash Prom on the processor board.

6.3.1.2	Upgrading Firmware from Floppy

In order to update the scope firmware from floppy, two disks are required.

The first contains a 93XX36XX.bin file in the Lecroy_P directory.

The second contains a 93XX36XX.fla file in the Lecroy_P directory

§ Insert the first disk and cycle power to the scope. Once re-booted enter the Utilities, 	  	  Special Modes, Firmware Update menu.

§ Change the Update From control to Floppy

§ Insert the second floppy disk and press Update Program.

The Software is then downloaded to the Flash Prom.

## 6.3.2	Changing Software Options

The software option selection GAL is located on the processor board at location A49.

Insert or replace the GAL to select new options.

Make sure that the orientation notch is correctly aligned with the PCB.

## 6.3.3	Software Option Selection GAL

The following software options are available : ( see section 2 )

WP01          	Advanced Math Firmware

WP02          	Basic FFT Firmware

WP03	Parameter Distribution Analysis Firmware

DDM	Disk Drive Measurements

PRML	Partial Response Maximum Likelihood

ORM	Optical Recording Measurement Package

MC01	Memory Card Reader

| PAL | MC01 | ORM | PRML | DDM | WP03 | WP02 | WP01 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| KEY001-A |  |  |  |  |  |  | WP01 |
| KEY002-A |  |  |  |  |  | WP02 |  |
| KEY003-A |  |  |  |  |  | WP02 | WP01 |
| KEY004-A |  |  |  |  | WP03 |  |  |
| KEY005-A |  |  |  |  | WP03 |  | WP01 |
| KEY006-A |  |  |  |  | WP03 | WP02 |  |
| KEY007-A |  |  |  |  | WP03 | WP02 | WP01 |
| KEY008-A |  |  |  | DDM |  |  |  |
| KEY009-A |  |  |  | DDM |  |  | WP01 |
| KEY00A-A |  |  |  | DDM |  | WP02 |  |
| KEY00B-A |  |  |  | DDM |  | WP02 | WP01 |
| KEY00C-A |  |  |  | DDM | WP03 |  |  |
| KEY00D-A |  |  |  | DDM | WP03 |  | WP01 |
| KEY00E-A |  |  |  | DDM | WP03 | WP02 |  |
| KEY00F-A |  |  |  | DDM | WP03 | WP02 | WP01 |
| KEY020-A |  |  | PRML |  |  |  |  |
| KEY021-A |  |  | PRML |  |  |  | WP01 |
| KEY022-A |  |  | PRML |  |  | WP02 |  |
| KEY023-A |  |  | PRML |  |  | WP02 | WP01 |
| KEY024-A |  |  | PRML |  | WP03 |  |  |
| KEY025-A |  |  | PRML |  | WP03 |  | WP01 |
| KEY026-A |  |  | PRML |  | WP03 | WP02 |  |
| KEY027-A |  |  | PRML |  | WP03 | WP02 | WP01 |
| KEY028-A |  |  | PRML | DDM |  |  |  |
| KEY029-A |  |  | PRML | DDM |  |  | WP01 |
| KEY02A-A |  |  | PRML | DDM |  | WP02 |  |
| KEY02B-A |  |  | PRML | DDM |  | WP02 | WP01 |
| KEY02C-A |  |  | PRML | DDM | WP03 |  |  |
| KEY02D-A |  |  | PRML | DDM | WP03 |  | WP01 |
| KEY02E-A |  |  | PRML | DDM | WP03 | WP02 |  |
| KEY02F-A |  |  | PRML | DDM | WP03 | WP02 | WP01 |
| KEY040-A |  | ORM |  |  |  |  |  |
| KEY041-A |  | ORM |  |  |  |  | WP01 |
| KEY042-A |  | ORM |  |  |  | WP02 |  |
| KEY043-A |  | ORM |  |  |  | WP02 | WP01 |
| KEY044-A |  | ORM |  |  | WP03 |  |  |
| KEY045-A |  | ORM |  |  | WP03 |  | WP01 |
| KEY046-A |  | ORM |  |  | WP03 | WP02 |  |
| KEY047-A |  | ORM |  |  | WP03 | WP02 | WP01 |
| KEY048-A |  | ORM |  | DDM |  |  |  |
| KEY049-A |  | ORM |  | DDM |  |  | WP01 |
| KEY04A-A |  | ORM |  | DDM |  | WP02 |  |
| KEY04B-A |  | ORM |  | DDM |  | WP02 | WP01 |
| KEY04C-A |  | ORM |  | DDM | WP03 |  |  |
| KEY04D-A |  | ORM |  | DDM | WP03 |  | WP01 |
| KEY04E-A |  | ORM |  | DDM | WP03 | WP02 |  |
| KEY04F-A |  | ORM |  | DDM | WP03 | WP02 | WP01 |
| KEY060-A |  | ORM | PRML |  |  |  |  |
| PAL | MC01 | ORM | PRML | DDM | WP03 | WP02 | WP01 |
| KEY061-A |  | ORM | PRML |  |  |  | WP01 |
| KEY062-A |  | ORM | PRML |  |  | WP02 |  |
| KEY063-A |  | ORM | PRML |  |  | WP02 | WP01 |
| KEY064-A |  | ORM | PRML |  | WP03 |  |  |
| KEY065-A |  | ORM | PRML |  | WP03 |  | WP01 |
| KEY066-A |  | ORM | PRML |  | WP03 | WP02 |  |
| KEY067-A |  | ORM | PRML |  | WP03 | WP02 | WP01 |
| KEY068-A |  | ORM | PRML | DDM |  |  |  |
| KEY069-A |  | ORM | PRML | DDM |  |  | WP01 |
| KEY06A-A |  | ORM | PRML | DDM |  | WP02 |  |
| KEY06B-A |  | ORM | PRML | DDM |  | WP02 | WP01 |
| KEY06C-A |  | ORM | PRML | DDM | WP03 |  |  |
| KEY06D-A |  | ORM | PRML | DDM | WP03 |  | WP01 |
| KEY06E-A |  | ORM | PRML | DDM | WP03 | WP02 |  |
| KEY06F-A |  | ORM | PRML | DDM | WP03 | WP02 | WP01 |
| KEY200-A | MC01 |  |  |  |  |  |  |
| KEY201-A | MC01 |  |  |  |  |  | WP01 |
| KEY202-A | MC01 |  |  |  |  | WP02 |  |
| KEY203-A | MC01 |  |  |  |  | WP02 | WP01 |
| KEY204-A | MC01 |  |  |  | WP03 |  |  |
| KEY205-A | MC01 |  |  |  | WP03 |  | WP01 |
| KEY206-A | MC01 |  |  |  | WP03 | WP02 |  |
| KEY207-A | MC01 |  |  |  | WP03 | WP02 | WP01 |
| KEY208-A | MC01 |  |  | DDM |  |  |  |
| KEY209-A | MC01 |  |  | DDM |  |  | WP01 |
| KEY20A-A | MC01 |  |  | DDM |  | WP02 |  |
| KEY20B-A | MC01 |  |  | DDM |  | WP02 | WP01 |
| KEY20C-A | MC01 |  |  | DDM | WP03 |  |  |
| KEY20D-A | MC01 |  |  | DDM | WP03 |  | WP01 |
| KEY20E-A | MC01 |  |  | DDM | WP03 | WP02 |  |
| KEY20F-A | MC01 |  |  | DDM | WP03 | WP02 | WP01 |
| KEY220-A | MC01 |  | PRML |  |  |  |  |
| KEY221-A | MC01 |  | PRML |  |  |  | WP01 |
| KEY222-A | MC01 |  | PRML |  |  | WP02 |  |
| KEY223-A | MC01 |  | PRML |  |  | WP02 | WP01 |
| KEY224-A | MC01 |  | PRML |  | WP03 |  |  |
| KEY225-A | MC01 |  | PRML |  | WP03 |  | WP01 |
| KEY226-A | MC01 |  | PRML |  | WP03 | WP02 |  |
| KEY227-A | MC01 |  | PRML |  | WP03 | WP02 | WP01 |
| KEY228-A | MC01 |  | PRML | DDM |  |  |  |
| KEY229-A | MC01 |  | PRML | DDM |  |  | WP01 |
| KEY22A-A | MC01 |  | PRML | DDM |  | WP02 |  |
| KEY22B-A | MC01 |  | PRML | DDM |  | WP02 | WP01 |
| KEY22C-A | MC01 |  | PRML | DDM | WP03 |  |  |
| KEY22D-A | MC01 |  | PRML | DDM | WP03 |  | WP01 |
| KEY22E-A | MC01 |  | PRML | DDM | WP03 | WP02 |  |
| KEY22F-A | MC01 |  | PRML | DDM | WP03 | WP02 | WP01 |
| KEY240-A | MC01 | ORM |  |  |  |  |  |
| KEY241-A | MC01 | ORM |  |  |  |  | WP01 |
| KEY242-A | MC01 | ORM |  |  |  | WP02 |  |
| PAL | MC01 | ORM | PRML | DDM | WP03 | WP02 | WP01 |
| KEY243-A | MC01 | ORM |  |  |  | WP02 | WP01 |
| KEY244-A | MC01 | ORM |  |  | WP03 |  |  |
| KEY245-A | MC01 | ORM |  |  | WP03 |  | WP01 |
| KEY246-A | MC01 | ORM |  |  | WP03 | WP02 |  |
| KEY247-A | MC01 | ORM |  |  | WP03 | WP02 | WP01 |
| KEY248-A | MC01 | ORM |  | DDM |  |  |  |
| KEY249-A | MC01 | ORM |  | DDM |  |  | WP01 |
| KEY24A-A | MC01 | ORM |  | DDM |  | WP02 |  |
| KEY24B-A | MC01 | ORM |  | DDM |  | WP02 | WP01 |
| KEY24C-A | MC01 | ORM |  | DDM | WP03 |  |  |
| KEY24D-A | MC01 | ORM |  | DDM | WP03 |  | WP01 |
| KEY24E-A | MC01 | ORM |  | DDM | WP03 | WP02 |  |
| KEY24F-A | MC01 | ORM |  | DDM | WP03 | WP02 | WP01 |
| KEY260-A | MC01 | ORM | PRML |  |  |  |  |
| KEY261-A | MC01 | ORM | PRML |  |  |  | WP01 |
| KEY262-A | MC01 | ORM | PRML |  |  | WP02 |  |
| KEY263-A | MC01 | ORM | PRML |  |  | WP02 | WP01 |
| KEY264-A | MC01 | ORM | PRML |  | WP03 |  |  |
| KEY265-A | MC01 | ORM | PRML |  | WP03 |  | WP01 |
| KEY266-A | MC01 | ORM | PRML |  | WP03 | WP02 |  |
| KEY267-A | MC01 | ORM | PRML |  | WP03 | WP02 | WP01 |
| KEY268-A | MC01 | ORM | PRML | DDM |  |  |  |
| KEY269-A | MC01 | ORM | PRML | DDM |  |  | WP01 |
| KEY26A-A | MC01 | ORM | PRML | DDM |  | WP02 |  |
| KEY26B-A | MC01 | ORM | PRML | DDM |  | WP02 | WP01 |
| KEY26C-A | MC01 | ORM | PRML | DDM | WP03 |  |  |
| KEY26D-A | MC01 | ORM | PRML | DDM | WP03 |  | WP01 |
| KEY26E-A | MC01 | ORM | PRML | DDM | WP03 | WP02 |  |
| KEY26F-A | MC01 | ORM | PRML | DDM | WP03 | WP02 | WP01 |

## 6.3.4	Processor Board Exchange Procedure

The replacement board is supplied without any options. Therefore the existing GAL

at Loc. A49, must be transferred from the faulty board to the new board.

After upgrading firmware or changing the software option, check that the scope boots 	correctly.

Then check in the system summary, by using the show status button on the front panel, 	the software version, software options and serial number.

The serial number of the 9384 oscilloscope is loaded in the real time clock memory 	which is battery backed up. If it becomes necessary to replace the processor board, the 	serial number must be loaded in the memory of the new board by using LeCroy program 	" LeCalsoft " under GPIB remote control.

To run " LeCalsoft " type SKP.exe, in the main menu type S, and follow the 				instructions, use five digits to enter the serial number ( i.e. 01732 ).

## 6.4	Equipment and Spare Parts Recommended for Service

## 6.4.1	Equipment

The following equipment is needed to provide the technician access to the 9384

subassemblies during repair and calibration

See also Performance Verification section 5.

| Instrument | Qty | Specifications | Recommended |
| --- | --- | --- | --- |
| Signal Generator ( sine wave ) | 1 | Frequency : .5 MHz to 2 GHz Accuracy : 1ppm Amplitude : 5 V peak to peak | HP8648B or equivalent |
| Digital Multimeter | 1 | 5 digits | Keithley 2000 |
| Fast pulser | 1 | Rise time < 500 psec | LeCroy 4969 |
| Cable | 1 | BNC, 50 W, length 20 cm & 100 cm ( 6.87 & 39.37 inches ) | Suhner |

## 6.4.2	Spare Parts

In order to make the repair of 9384 oscilloscope at board level, a minimum stock of 	boards is at least one each:

	F9302-1-8	:	Processor board

	9384-3	:	Main board

	9384M-2	:	1 M Memory board

	9384MEM-2	:	2 M Memory board

	F9300-4	:	GPIB/RS232 interface

	F9354-5	:	Front panel with keyboard

	93XX-Display	:	Raster monitor kit ( deflection, video, yoke & CRT )

	PS9384	:	Power supply

If the unit is equipped with the 93XX-FD01 option :

§ F9300-6	:	Floppy, Graphic printer, Centronics Interface

§ 335023203	:	Floppy disk drive

If the unit is equipped with the 93XX-GP01 option :

§ F9300-6	:	Graphic printer, Floppy, Centronics Interface

§ F9300-7	:	Graphic printer controller

§ 334000832	:	LPT5446 Seiko Graphic printer

If the unit is equipped with the 93XX-HD01 option :

F9300-8	:	Hard disk Interface

## HDD02:	Hard disk drive

## 6.5	Troubleshooting and Flow Charts

Most procedures in this section will allow troubleshooting down to the BOARD LEVEL.

Defective circuit boards will be repaired or exchanged by the regional LeCroy service 	office or the local representative (see section 1.4).

## 6.5.1	Introduction

The troubleshooting information contained in this section is intended for use by qualified 	personnel having a basic understanding of electronics (analog and digital). In order to 	simplify servicing and minimize downtime, the following list of possible symptoms, 	likely causes, and troubleshooting steps have been prepared.

The first step in troubleshooting is to check for obvious items like blown fuses.

The power supply is the next item to check before proceeding to more detailed 	troubleshooting, since noise or low power supply voltages can cause a variety of digital 	and analog problems.

## 6.5.2	Line Voltage Autoranging

The 9384 oscilloscope operates from a 115 V (90 to 130 V) or 220 V

(180 to 260 V) normal power source at 47 Hz to 63Hz.

No voltage selection is required since the instrument automatically adapts to the line 	voltage which is present.

The instrument operates at line frequencies up to 440 Hz.

## 6.5.3	Initial Troubleshooting Chart

## 6.5.4	No Power Supply

## 6.5.4.1	Line Fuses Replacement

The power supply of the oscilloscope is protected against short circuits and overload by 	means of two 6.3 A / 250 V fuses located above the main plugs.

### W A R N I N G

Disconnect the instrument from the power line and from other equipment before 	replacing fuses.

To replace line fuses, proceed as follow :

§ Turn off the power and disconnect the line cord from the instrument

§ Open the fuse box by inserting a small flat screwdriver under the plastic cover and 		   remove the fuse carrier from the holder

§ Remove the 6.3 amp fuse and replace it with the proper type:

6.3 amp/250 V,  normal blowing.

## LeCroy part number: 433 162 630

## 6.5.5	No Display

## 6.5.6	Abnormal Image On Screen

6.5.7	Front Panel Controls Do Not Operate

## 6.5.8	No Remote Control GPIB and RS232

## 6.5.9	Performance Verification Fails

6.5.10	Floppy Disk Drive Fails

## 6.5.11	Graphic Printer Fails

## 6.5.13	Centronics Fails

## 6.5.14	Hard Disk Fails

## 6.6		Calibration Procedures

The following section includes the adjustments required for the power supply and 	display. It is recommended that they be verified at one year intervals.

## 6.6.1	PS9384 Power Supply Calibration

The five voltages are adjustable by ± 5% of the nominal value.

The reference for the measurements are the pins on top of connector J31 located on the 	main board F9384-3.

For the power supply calibration proceed as follow:

§ Turn off the power

§ Remove the top cover (6.2.1)

§ Remove the front frame assembly (6.2.9) and put it to the right of the unit.

§ By using two extension cables, reconnect the processor board to the front panel (J4) 	  	  and to the deflection board (J6).

§ Once the top cover is removed and the front panel is disassembled from the scope, extra 	   cooling of the main board is required. It's mandatory to disconnect the existing Fan 	   	   from connector J33, located on F9384-3 card, and to use a Fan with the air flow 	   	   oriented to the front end section of the board.

§ The front frame assembly is now reconnected to the processor through the extension 	  	   cables.

§ Turn on the power, set the scope to Auto Trigger, and perform the adjustments to get on 	   J31 :

| 12 | 24 |
| --- | --- |
| 11 | 23 |
| 10 | 22 |
| 9 | 21 |
| 8 | 20 |
| 7 | 19 |
| 6 | 18 |
| 5 | 17 |
| 4 | 16 |
| 3 | 15 |
| 2 | 14 |
| 1 | 13 |
| Front |  |

Pin 15		:	+15.0 V  ( Min = + 15.4 V, Max. = + 15.6 V )

Pin 14		:	-15.0 V ( Min = - 15.6 V, Max. = - 15.4 V )

Pin 13		:	-6.0 V ( Min = - 7.10 V, Max. = - 6.0 V )

Pin 10,11,12		:	+3.3 V  ( Min = + 3.22 V, Max. = + 3.38 V )

Pin 9, 20, 21		:	+5.18 V  ( Min = + 5.13 V, Max. = + 5.23 V )

Pin 6,7,8,19,23,24	:	Ground

Pin 4,5,17,18	:	-5.2 V ( Min = - 5.15 V, Max. = + 5.15 V )

Pin  2,3,16		:	-2.1 V ( Min = - 2.15 V, Max. = -2.05 V )

Pin 1		:	+6.0 V ( Min = + 6.0 V, Max. = + 7.1 V )

The four potentiometers R108, R206, R306, and R407 are accessible from the front  	through holes in the PS9384 power supply chassis.  In order to adjust the -2v supply, the 	PS9384 must be removed from the unit.  The adjustment hole is located on the rear side 	of the power supply.

§ Turn the potentiometer clockwise to increase the tension or counterclockwise to 	  	   decrease the voltage. When the adjustment is done, stop the acquisition by depressing 	   the stop trigger push button, and verify that there is no large difference on the + 5.10 V,  	   typically less than 30 mv.

## Figure 6.9 : Power Supply

## 6.6.2	93XX-Display Adjustment Procedure

## 6.6.2.1	Introduction

There is a total of 12 potentiometers or variable coils to adjust the deflection and video 	board.

Video: (2 adjustments)

§ Threshold	:    Level of the video board.

§ Gain	:    Intensity of the screen.

Video board component side

Deflection	:	(10 adjustments)

§ Vosc	:	Frequency of the vertical oscillator.

§ Slope	:	Speed of the horizontal ramp.

§ Focus	:	Focus  of the screen.

§ Cut off	:	Cathode ray tube cut off.

§ Quiescent	:	Standby current of the horizontal deflection amplifier.

§ H Linearity	:	Horizontal linearity.

§ H Size	:	Horizontal size (Max. 165mm).

§ H Offset	:	Horizontal position.

§ V Size	:	Vertical size (Max. 120mm).

§ V Offset	:	Vertical position.

Deflection board component side

## 6.6.2.2  Coarse Adjustment

§ Depress display button.

§ Set W’form + text intensity to 0%.

§ Set grid intensity to 0%

§ Turn fully clockwise the intensity potentiometer on the video board.

§ On the video board connect a digital multimeter on test point : W303

§ Adjust threshold potentiometer to get 2 V ± 0.1 V on W303.

§ Set W'form intensity to 100%.

§ Set grid intensity to 60%.

§ Adjust H-size, H-offset, V-size, V-offset to center the image in the screen.

The vertical position should be adjusted to get the push buttons of the front panel in 	  	  front of the software menus, use the utilities set up.

The small magnets mounted on the deflection yoke influence the vertical position.

§ Turn the quiescent potentiometer clockwise until the default of the horizontal lines just 	   disappears from the vertical center of the screen.

§ Increase the cut off until a vertical line appears on the right side of the screen.

§ Adjust the slope potentiometer to get 5 mm gap between the highlighted vertical line 	  	  and the right border of the selection menus.

§ Adjust H-linearity to get the best linearity.

## 6.6.2.3  Fine Adjustment

The final adjustment of the intensity, cut off, and focus must be made in a dark room.

§ Set W'form intensity to 30%.

§ Set grid intensity to 0%.

§ Adjust the cut off potentiometer until the highlighted vertical line disappears from the 	   right side of the screen.

§ Set W'form intensity to 20%.

§ Display four traces.

§ On the video board adjust the gain potentiometer (intensity) in order to get the text

just readable.

§ Set W'form + text intensity to 70%.

§ Set grid intensity to 30%

§ Adjust the focus (usually fully clockwise) for most uniform focus

over the entire screen.

§ In a standard luminosity environment set W'form + text to 90%, and

grid intensity to 60%.

§ Verify the intensity, focus, and contrast adjustment, for best

definition of the displayed text.

### C A U T I O N

Never change the Vosc calibration.