Earthy to Floating Processors messaging system

The basis of the Earthy and Floating Processors interface is a simple serial link and special messaging protocol. Serial link have a two selectable clocks: 76800 bauds for normal operation and 300 bauds for debug purposes.

Start bit

Stop bit

8 data bits

130 us

…

13 us

~ 50 ms

…

Table 1 – Single message structure

| Byte # | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Function | 81h | CRC | Range | Mode | Nines | State | Cmd | 8Dh |

| Byte # | Function | Description |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 1 | 81h | Begin of the message flag / Message type |  |  |  |
| 2 | CRC | Message checksum = 0FFh – sum(byte[3:8]) |  |  |  |
| 3 | Range | DMM Range selection. For DCV: |  |  |  |
|  |  | 0 = Auto | 1 = 0.1 V | 2 = 1.0 V | 3 = 10 V |
|  |  | 4 = 100 V | 5 = 1000 V |  |  |
| 4 | Mode | DMM Mode selection: |  |  |  |
|  |  | 0 = VDC | 1 = VAC | 2 = Ohms | 3 = VAC + VDC |
|  |  | 4 = Test0 | 5 = Test10 | 6 = TestK | 7 = TestAC ? |
|  |  | 8 = Reference | 9 = [1]+filt | 0Ah = [3] + filt | 0Bh = True Ohms |
| 5 | Nines | Integration time / Digits selection: |  |  |  |
|  |  | 3 = 3x9 | 4 = 4x9 | 5 = 5x9 | 6 = 6x9 |
|  |  | 7 = 7x9 | 8 = 8x9 | * 7x9 temporarily used while Null |  |
| 6 | State | Internal state flags (0/1 = Off/On): |  |  |  |
|  |  | bit 0 = Ratio measurement |  | bit 1 = Drift correction (~15 min.) |  |
| 7 | Cmd | Special commands: |  |  |  |
|  |  | 0 = Disable tracking |  | 1 = One sample trigger |  |
|  |  | 2 = Enable tracking |  | 4 = Force drift correction now |  |
| 8 | 8Dh | End of String (EOS) / Acknowledge flag |  |  |  |

Besides standard messages there are another message types with different structures. For example:

8Dh – Floating to Earthy acknowledge;

82h, 71h, 8Dh – DUMP NVM command;

C0h, 8Dh – HISTORY command.