Earthy Logic I/O Map

| Base address, hex | Devices |
| --- | --- |
| 5C00 | IC27 –> IC16 ? |
| 5800 | GPIB Address Switches |
| 5400 | PIA: Minate interface, NVClock, Beeper |
| 5000 | R/WQ, Floating UART |
| 4C00 | PIA: Keyboard, NVClock, Calibration key |
| 4800 | RS232 |
| 4400 | GPIB I/O |
| 4000 | DMA Select |

Floating Logic Memory Map

| Address range, hex | Devices |  |
| --- | --- | --- |
| E000-FFFF | ROM |  |
| C000-DFFF | Free space (no decoding) |  |
| A000-A3FF, A400-A7FF | NVM pages |  |
| 8000-9FFF | Free space (no decoding) |  |
| 6000-7FFF | I/O space |  |
| 4000-5FFF | GLUG LO - | Read word at 3FFF-4000 |
| 2000-3FFF | GLUG HI - / Reset |  |
| 0000-1FFF | RAM, CPU internal registers and ports |  |

Floating Logic I/O Map

| Port | 6000h (OUT LO-) | 6100h (OUT HI-) | 6200h (OUT AC-) | 6300 (OUT TC-) |
| --- | --- | --- | --- | --- |
| bit 7 | RATIO Control | RLD7 (protection) | AC | FRSS (6302h) |
| bit 6 | I/P AMP Control | – | AC+DC | FAS (6304h) |
| bit 5 | RATIO HI Control | RLD501 (TestK) | AC Filter | TC6 |
| bit 4 | Test0 | RLD6 (in. Div100) | TestAC | TC5 |
| bit 3 | 10 uA Control | RLD5 (in. InpAmp) | AC ? | TC4 |
| bit 2 | x100 | RLD4 (Div100) | 1 | TC3 |
| bit 1 | x10 | RLD2 (Test10) | 100 | TC2 |
| bit 0 | x1 | RLD1 (Ohms) | 1000 | TC1 |

Floating Processor Ports

| Port 2 (Adress 0003h) |  |  |  |  | Port 1 (Adress 0002h) |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Function | Dir. | Name | bit |  | bit | Name | Dir. | Function |
| 400 Hz | < | P20 | 0 |  | 7 | P17 | < | Ext. sw. (Front Panel, SCI bauds) |
| MPU type | > | P21 | 1 |  | 6 | P16 | < | MAIN frequency detect |
| not 60 Hz | < | P22 | 2 |  | 5 | P15 | > | I/O Bus: NVM C1; TC6 |
| SCI E to F | > | P23 | 3 |  | 4 | P14 | > | I/O Bus: NVM C0; TC5 |
| SCI F to E | < | P24 | 4 |  | 3 | P13 | <> | I/O Bus: NVM D3; TC4 |
|  |  |  |  |  | 2 | P12 | <> | I/O Bus: NVM D2; TC3 |
|  |  |  |  |  | 1 | P11 | <> | I/O Bus: NVM D1; TC2 |
|  |  |  |  |  | 0 | P10 | <> | I/O Bus: NVM D0; TC1 |