NEUTRIK RT-1M Audio Test System (VISA)

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for NEUTRIK RT-1M Audio Test System (VISA).

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

________________________________________________________________________________

Error and Status Information:

Error codes and warnings are returned as the return value of each instrument

driver function. A program should examine this value after each call to an

instrument driver function to determine if any error occurred.  The general

meaning of the return value is as follows:

Value                   Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

Possible error codes and their meanings are listed with the corresponding

instrument driver function.

________________________________________________________________________________

How To Use This Document:

This document is intended to be used as a programming reference manual.

It describes each function in the NEUTRIK RT-1M Audio Test System (VISA)

instrument driver. The functions are listed in alphabetical order with a

description of the function, C syntax of the function, a description of each

parameter, and possible error codes.

________________________________________________________________________________

Function Tree Layout:

NEUTRIK RT-1M Audio Test System (VISA) Function Name:

Initialize                                  ntrt1m_init

Application Functions

Simple Measurement                       ntrt1m_simpleMeas

Configuration Functions

Configure Input Mode                     ntrt1m_configInputMode

Configure Input Range                    ntrt1m_configInputRange

Configure Output Mode                    ntrt1m_configOutputMode

Configure Output Level                   ntrt1m_configOutputLevel

Configure Measurement Unit               ntrt1m_configMeasUnit

Configure Multitone                      ntrt1m_configMtone

Define Multitone Signal                  ntrt1m_defineMtoneSignal

Action/Status Functions

Start Active Multitone                   ntrt1m_startActiveMtone

Continuous Active Multitone              ntrt1m_contActiveMtone

Data Functions

Read Measurement                         ntrt1m_readMeas

Utility Functions

Write To Instrument                      ntrt1m_writeInstrData

Read Instrument Data                     ntrt1m_readInstrData

Reset                                    ntrt1m_reset

Self-Test                                ntrt1m_selfTest

Error-Query                              ntrt1m_errorQuery

Error Message                            ntrt1m_errorMessage

Revision Query                           ntrt1m_revisionQuery

Close                                       ntrt1m_close

________________________________________________________________________________

NEUTRIK RT-1M Audio Test System (VISA)

This instrument module provides programming support for the <Instrument

Name>.  The module is divided into the following functions:

Functions/Classes:

(1) Initialize:

This function initializes the instrument and sets it to a default

configuration.

(2) Application Functions: (Class)

This class contains high-level, test and measurement routines.  These

examples call other instrument driver functions to configure, start, and

read from the instrument.

(3) Configuration Functions: (Class)

This class of functions configures the instrument by setting acquisition

and system configuration parameters.

(4) Action/Status Functions: (Class)

This class of functions begins or terminates an acquisition. It also

provides functions which allow the user to determine the current status

of the instrument.

(5) Data Functions: (Class)

This class of functions transfers data to or from the instrument.

(6) Utility Functions: (Class)

This class of functions provides lower level functions to communicate

with the instrument, and change instrument parameters.

(7) Close:

This function takes the instrument offline.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

ntrt1m_close

ViStatus ntrt1m_close (ViSession instrumentHandle);

Purpose

This function performs the following operations:

viClose (instrSession) and viClose (rmSession).

Notes:

(1) The instrument must be reinitialized to use it again.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_configInputMode

ViStatus ntrt1m_configInputMode (ViSession instrumentHandle,

ViBoolean inputConnection,

ViBoolean CH1InternalLink,

ViBoolean CH2InternalLink,

ViInt16 synchronizationMode,

ViInt16 softwareWeightingFilter,

ViBoolean deemphasisFilter);

Purpose

This function selects either front or back panel connections for input,

links the selected output to the input of the analyzer internally,

selects source for synchronization on the expected multitone signal, and

enables one of the software filters.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

inputConnection

Variable Type       ViBoolean

This control selects the front or back panel connection for input.

Valid Range:

VI_OFF (0) - Back Panel Connection

VI_ON  (1) - Front Panel Connection   (Default)

CH1InternalLink

Variable Type       ViBoolean

This control links the selected output to the intput of the analyzer

internally.

Valid Range:

VI_OFF (0) - Link Off   (Default)

VI_ON  (1) - Link On

CH2InternalLink

Variable Type       ViBoolean

This control links the selected output to the intput of the analyzer

internally.

Valid Range:

VI_OFF (0) - Link Off   (Default)

VI_ON  (1) - Link On

synchronizationMode

Variable Type       ViInt16

This control selects the synchronization source.

Valid Range:

0 - Internal w/ header    (Default)

1 - Internal w/o header

2 - External w/ header

3 - External w/o header

Notes:

To select an external source, minimum one external measurement is

required beforehand.

softwareWeightingFilter

Variable Type       ViInt16

This control selects the software weighting filter for both channels.

Valid Range:

0 - OFF           (Default)

1 - C-Weighting

2 - CCITT

deemphasisFilter

Variable Type       ViBoolean

This control enables 750 microseconds Deemphasis filter.

Valid Range:

VI_OFF (0) - Disabled   (Default)

VI_ON  (1) - Enabled

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Input Connection) out of range.

BFFC0003  Parameter 3 (CH1 Internal Link) out of range.

BFFC0004  Parameter 4 (CH2 Internal Link) out of range.

BFFC0005  Parameter 5 (Synchronization Mode) out of range.

BFFC0006  Parameter 6 (Software Weighting Filter) out of range.

BFFC0007  Parameter 7 (Deemphasis Filter) out of range.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_configInputRange

ViStatus ntrt1m_configInputRange (ViSession instrumentHandle, ViInt16 input,

ViReal64 inputRange, ViBoolean inputUnit);

Purpose

This function sets the input range and preferred unit for the selected

channel.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

input

Variable Type       ViInt16

This control select the input channel to be configured.  If 'Channel

1 & 2' is selected, the same setting will be applied to both

channels.

Valid Range:

0 - Channel 1 and 2   (Default)

1 - Channel 1 only

2 - Channel 2 only

inputRange

Variable Type       ViReal64

This control sets input range of the selected channel(s).

Valid Range:   Dependent upon the selected unit:

-60   to +20 dBVp

0.001 to  10 Vp

Default value:  0 dBVp

inputUnit

Variable Type       ViBoolean

This control selects the input unit for the selected channel(s).

Valid Range:

VI_OFF (0) - dBVp (Default)

VI_ON  (1) - Vp

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Input) out of range.

BFFC0003  Parameter 3 (Input Range) out of range.

BFFC0004  Parameter 4 (Input Unit) out of range.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_configMeasUnit

ViStatus ntrt1m_configMeasUnit (ViSession instrumentHandle, ViInt16 configure,

ViInt16 levelUnit, ViBoolean distortionUnit,

ViBoolean noiseUnit, ViBoolean phaseUnit,

ViReal64 phaseScale, ViBoolean crosstalk);

Purpose

This function selects the preferred units of level, distortion, noise,

and crosstalk for the selected channel.  This function also selects the

preferred unit and sets the scaling of phase for both channels.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

configure

Variable Type       ViInt16

This control select the channel to be configured.  If 'Channel 1 & 2'

is selected, the same setting will be applied to both channels.

Valid Range:

0 - Channel 1 and 2   (Default)

1 - Channel 1 only

2 - Channel 2 only

levelUnit

Variable Type       ViInt16

This control selects the unit of level for the selected channel(s).

Valid Range:

0 - dBVp   (Default)

1 - dBV

2 - Vp

3 - V

distortionUnit

Variable Type       ViBoolean

This control selects the unit of distortion for the selected

channel(s).

Valid Range:

VI_OFF (0) - dBV   (Default)

VI_ON  (1) - V

noiseUnit

Variable Type       ViBoolean

This control selects the unit of noise for the selected channel(s).

Valid Range:

VI_OFF (0) - dBV   (Default)

VI_ON  (1) - V

phaseUnit

Variable Type       ViBoolean

This control selects the unit of phase for both channel(s).

Valid Range:

VI_OFF (0) - dBV   (Default)

VI_ON  (1) - V

Notes:

The same phase unit will be applied to both channels, even if only

one channel is selected to configure measurement units.

phaseScale

Variable Type       ViReal64

This control sets the scaling of the phase-plot for both channels.

Valid Range:   Dependent upon the selected unit:

-6.283185308 to 0.0 radian

-360         to 0.0 degree

Default Value:  0.0

Notes:

Setting this control to -180 (degree) sets the scale to -180 to +180

degrees.

crosstalk

Variable Type       ViBoolean

This control selects the unit of crosstalk measurement for the

selected channel(s).

Valid Range:

VI_OFF (0) - radian   (Default)

VI_ON  (1) - degree

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Configure) out of range.

BFFC0003  Parameter 3 (Level Unit) out of range.

BFFC0004  Parameter 4 (Distortion Unit) out of range.

BFFC0005  Parameter 5 (Noise Unit) out of range.

BFFC0006  Parameter 6 (Phase Unit) out of range.

BFFC0007  Parameter 7 (Phase Scale) out of range.

BFFC0008  Parameter 8 (Crosstalk) out of range.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_configMtone

ViStatus ntrt1m_configMtone (ViSession instrumentHandle, ViInt16 activeSignal,

ViReal64 pretriggerLengthms,

ViReal64 multitoneBurstLengthms);

Purpose

This function selects the active signal to be used for transmission from

1 of 4 memory locations, and sets the duration of the pre-trigger signal

and the multitone burst in milliseconds.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

activeSignal

Variable Type       ViInt16

This control selects the active multitone signal.

Valid Range:

1 - Signal in Memory 1   (Default)

2 - Signal in Memory 2

3 - Signal in Memory 3

4 - Signal in Memory 4

pretriggerLengthms

Variable Type       ViReal64

This control sets the pre-trigger signal duration in milliseconds for

the active signal.

Valid Range:

0 to 30000 ms

Default Value:  0 ms

multitoneBurstLengthms

Variable Type       ViReal64

This control sets the multitone burst duration in milliseconds for

the active signal.

Valid Range:

0 to 30000 ms

Default Value:  0 ms

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Active Signal) out of range.

BFFC0003  Parameter 3 (Pretrigger Length) out of range.

BFFC0004  Parameter 4 (Multitone Burst Length) out of range.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_configOutputLevel

ViStatus ntrt1m_configOutputLevel (ViSession instrumentHandle, ViInt16 output,

ViReal64 outputLevel, ViBoolean outputUnit);

Purpose

This function sets the output level and preferred unit for the selected

channel.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

output

Variable Type       ViInt16

This control select the output channel to be configured.  If 'Channel

1 & 2' is selected, the same setting will be applied to both

channels.

Valid Range:

0 - Channel 1 and 2   (Default)

1 - Channel 1 only

2 - Channel 2 only

outputLevel

Variable Type       ViReal64

This control sets the output level of the selected channel(s).

Valid Range:   Dependent upon the selected unit:

-60   to +20 dBVp

0.001 to  10 Vp

Default value:  0 dBVp

outputUnit

Variable Type       ViBoolean

This control selects the output unit for the selected channel(s).

Valid Range:

VI_OFF (0) - dBVp (Default)

VI_ON  (1) - Vp

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Output) out of range.

BFFC0003  Parameter 3 (Output Level) out of range.

BFFC0004  Parameter 4 (Output Unit) out of range.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_configOutputMode

ViStatus ntrt1m_configOutputMode (ViSession instrumentHandle,

ViBoolean floatGroundMode,

ViBoolean channel1Mute,

ViBoolean channel2Mute);

Purpose

This function selects either floating or grounded mode for both output

channels, and mutes the selected output channels.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

floatGroundMode

Variable Type       ViBoolean

This control selects either floating or grounded mode for both

channels.

Valid Range:

VI_OFF (0) - Ground Mode   (Default)

VI_ON  (1) - Float Mode

channel1Mute

Variable Type       ViBoolean

This control selects mute on or off for the channel 1.

Valid Range:

VI_OFF (0) - Mute Off   (Default)

VI_ON  (1) - Mute On

channel2Mute

Variable Type       ViBoolean

This control selects mute on or off for the channel 2.

Valid Range:

VI_OFF (0) - Mute Off   (Default)

VI_ON  (1) - Mute On

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Float Ground Mode) out of range.

BFFC0003  Parameter 3 (Channel 1 Mute) out of range.

BFFC0004  Parameter 4 (Channel 2 Mute) out of range.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_contActiveMtone

ViStatus ntrt1m_contActiveMtone (ViSession instrumentHandle);

Purpose

This function causes the generator to start with the active multitone

signal in an endless loop.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_defineMtoneSignal

ViStatus ntrt1m_defineMtoneSignal (ViSession instrumentHandle,

ViInt16 signalNumber, ViChar signalName[],

ViInt16 numberofSamples,

ViInt16 numberofBinsinCH1,

ViInt16 CH1BinFreqHz[],

ViReal64 CH1BinPhase[],

ViInt16 numberofBinsinCH2,

ViInt16 CH2BinFreqHz[],

ViReal64 CH2BinPhase[]);

Purpose

This function defines the user-specified multitone signal, and it stores

the signal to the selected memory location.  The following parameters

must be specified.

- Memory location and signal identification

- Signal bins / bandwidth

- Number of samples /blocklength

- Phases / Crest factor

Notes:

The signal definition procedure has to be performed only if a new

multitone signal has to be introduced.  Defined signals are kept in

RT-1M's memory even during power off.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

signalNumber

Variable Type       ViInt16

This control selects RT-1M's target memory.

Valid Range:

1 to 4

signalName

Variable Type       ViChar[]

This control sets user-defined signal name, enclosed in apostrophes

(').

Valid Range:

Up to 8 ASCII characters

numberofSamples

Variable Type       ViInt16

This control selects the number of samples (blocklength).

Valid Range:

0 - 512 Samples

1 - 1024 Samples

2 - 2048 Samples

3 - 4096 Samples

4 - 8192 Samples

Notes:

Blocklength     Min. burst duration     Typ. burst duration

(without header)        (with header)

512             154ms                   260ms

1024            284ms                   390ms

2048            344ms                   450ms

4096            684ms                   790ms

8192            854ms                   960ms

numberofBinsinCH1

Variable Type       ViInt16

This control sets the number of singnal bins to be transmitted in

channel 1.

Valid Range:

1 to 31

CH1BinFreqHz

Variable Type       ViInt16[]

This control contains an array of values that set the frequency for

each signal bin of channel 1.

Valid Range: for each array element

20 to 20000 Hz

Notes:

Bin-number (1 to 3413) is calculated according to the equation below.

Number of Samples

BinNumber = ------------------------ * Frequency

Internal Sampling Rate

where

Number of Samples = 512, 1024, 2048, 4096, or 8192

Frequency = 20 to 20000 Hz

Internal Sampling Rate = 48000 Hz

(1/48000 Hz = 21 microseconds of time interval)

CH1BinPhase

Variable Type       ViReal64[]

This control contains an array of values that set the phase of each

signal bin of channel 1.

Valid Range:

-3.141592654 to 3.141592654 rad

numberofBinsinCH2

Variable Type       ViInt16

This control sets the number of singnal bins to be transmitted in

channel 2.

Valid Range:

1 to 31

CH2BinFreqHz

Variable Type       ViInt16[]

This control contains an array of values that set the frequency for

each signal bin of channel 2.

Valid Range:

20 to 20000 Hz

Notes:

Bin-number (1 to 3413) is calculated according to the equation below.

Number of Samples

BinNumber = ------------------------ * Frequency

Internal Sampling Rate

where

Number of Samples = 512, 1024, 2048, 4096 or 8192

Frequency = 20 to 20000 Hz

Internal Sampling Rate = 48000 Hz

(1/48000 Hz = 21 microseconds of time interval)

CH2BinPhase

Variable Type       ViReal64[]

This control contains an array of values that set the phase of each

signal bin of channel 2.

Valid Range:

-3.141592654 to 3.141592654 rad

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Signal Number) out of range.

BFFC0004  Parameter 4 (Number of Samples) out of range.

BFFC0005  Parameter 5 (Number of Bins in CH1) out of range.

BFFC0008  Parameter 8 (Number of Bins in CH2) out of range.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_errorMessage

ViStatus ntrt1m_errorMessage (ViSession instrumentHandle, ViStatus errorCode,

ViChar errorMessage[]);

Purpose

This function takes the Status Code returned by the instrument driver

functions, interprets it and returns it as a user readable string.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

errorCode

Variable Type       ViStatus

This control accepts the Status Code returned from the instrument

driver functions.

Default Value: 0 - VI_SUCCESS

errorMessage

Variable Type       ViChar[]

This control returns the interpreted Status Code as a user readable

message string.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFF0085  VI_WARN_UNKNOWN_STATUS

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_errorQuery

ViStatus ntrt1m_errorQuery (ViSession instrumentHandle, ViInt32 *errorCode,

ViChar errorMessage[]);

Purpose

This function reads an error code from the instrument's error queue.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

errorCode

Variable Type       ViInt32 (passed by reference)

This control returns the error code read from the instrument's error

queue.

errorMessage

Variable Type       ViChar[]

This control returns the error message string read from the

instrument's error message queue.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_init

ViStatus ntrt1m_init (ViRsrc resourceName, ViBoolean IDQuery,

ViBoolean resetDevice, ViSession *instrumentHandle);

Purpose

This function performs the following initialization actions:

- Opens a session to the Default Resource Manager resource and a session

to the specified device using the interface and address specified in the

Resource_Name control.

- Performs an identification query on the Instrument.

- Resets the instrument to a known state.

- Sends initialization commands to the instrument that set any necessary

programmatic variables such as Headers Off, Short Command form, and Data

Transfer Binary to the state necessary for the operation of the

instrument driver.

- Returns an Instrument Handle which is used to differentiate between

different sessions of this instrument driver.

- Each time this function is invoked a Unique Session is opened.  It is

possible to have more than one session open for the same resource.

Notes:

(1) If this instrument does not support an ID Query, and the ID Query

control is set to "Do Query" then this function should return the Warning

Code 0x3FFC0101 - VI_WARN_NSUP_ID_QUERY.

(2) If this instrument does not support a Reset, and the Reset control is

set to "Reset Device" then this function should return the Warning Code

0x3FFC0102 - VI_WARN_NSUP_RESET.

Parameter List

resourceName

Variable Type       ViRsrc

This control specifies the interface and address of the device that

is to be initialized (Instrument Descriptor). The exact grammar to be

used in this control is shown in the note below.

Default Value:  "GPIB::3"

Notes:

(1) Based on the Instrument Descriptor, this operation establishes a

communication session with a device.  The grammar for the Instrument

Descriptor is shown below.  Optional parameters are shown in square

brackets ([]).

Interface   Grammar

------------------------------------------------------

GPIB        GPIB[board]::primary address[::secondary address]

[::INSTR]

The GPIB keyword is used with GPIB instruments.

The default value for optional parameters are shown below.

Optional Parameter          Default Value

-----------------------------------------

board                       0

secondary address           none - 31

IDQuery

Variable Type       ViBoolean

This control specifies if an ID Query is sent to the instrument

during the initialization procedure.

Valid Range:

VI_OFF (0) - Skip Query

VI_ON  (1) - Do Query (Default Value)

Notes:

(1) Under normal circumstances the ID Query ensures that the

instrument initialized is the type supported by this driver. However

circumstances may arise where it is undesirable to send an ID Query

to the instrument.  In those cases; set this control to "Skip Query"

and this function will initialize the selected interface, without

doing an ID Query.

resetDevice

Variable Type       ViBoolean

This control specifies if the instrument is to be reset to its

power-on settings during the initialization procedure.

Valid Range:

VI_OFF (0) - Don't Reset

VI_ON  (1) - Reset Device (Default Value)

Notes:

(1) If you do not want the instrument reset. Set this control to

"Don't Reset" while initializing the instrument.

instrumentHandle

Variable Type       ViSession (passed by reference)

This control returns an Instrument Handle that is used in all

subsequent function calls to differentiate between different sessions

of this instrument driver.

Notes:

(1) Each time this function is invoked a Unique Session is opened.

It is possible to have more than one session open for the same

resource.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (ID Query) out of range.

BFFC0003  Parameter 3 (Reset Device) out of range.

BFFC0011  Instrument returned invalid response to ID Query

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_readInstrData

ViStatus ntrt1m_readInstrData (ViSession instrumentHandle,

ViInt32 numberBytesToRead, ViChar readBuffer[],

ViInt32 *numBytesRead);

Purpose

This function reads data from the instrument's output buffer and returns

it to the specified variable in memory.

Notes:

(1) Because the instrument may return both numeric and text data in

response to queries this function returns the data in string format.

(2) If valid data is not available at the instrument's output buffer when

this function is called the instrument will hang up and the function will

not return until it times out.  If the time-out is disabled this function

will hang indefinitely and it may be necessary to reboot the computer to

break out.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

numberBytesToRead

Variable Type       ViInt32

The number of bytes specified by this control should be greater than

or equal to the number of bytes which are to be read from the

instrument. If the actual number of bytes to be read is greater than

the number this control specifies then multiple reads will be

required to empty the instrument's output buffer.

If the instrument's output buffer is not emptied the instrument may

return invalid data when future reads are performed.

Default Value:  50 (See NOTE 2)

Notes:

(1) If NO DATA is available at the instrument's output buffer when

this function is called the instrument will hang up and the function

will not return until it times out.  If the time-out is disabled,

this function will hang indefinitely and it may be necessary to

reboot the computer.

(2) If the number of bytes expected is greater than 50 the value of

this control may be increased. If the actual number of bytes read is

larger than the declared size of the read buffer a run-time error

will be generated.

readBuffer

Variable Type       ViChar[]

The incoming data from the instrument is placed into this variable.

Notes:

(1) Because the instrument may return both numeric and text data in

response to queries this function returns the data in string format.

(2) This function does not overwrite any old data left in the string

variable from the last time the function was called. String data in

LabWindows/CVI is terminated with an ASCII null(0x0) and string

manipulation functions will only recognize data before the ASCII

null.

(3) The declared size of the string variable must be greater than the

actual number of bytes read from the instrument if it is not a

run-time error will be generated.

numBytesRead

Variable Type       ViInt32 (passed by reference)

This variable contains the actual number of bytes read from the

instrument. This is the value which is returned by the read function.

Notes:

(1) If the actual number of bytes read is less than the number of

bytes specified in the Number Bytes To Read control then the output

buffer has probably been emptied. If the read function fails and the

number of bytes read is 0, the most probable cause for the failure is

there was no data available at the instrument's output buffer.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_readMeas

ViStatus ntrt1m_readMeas (ViSession instrumentHandle, ViInt16 selectChannel,

ViInt16 dataType, ViInt16 *numberofBinsQueried,

ViInt16 binNumber[], ViReal64 amplitude[],

ViChar unit[]);

Purpose

This function reads either the measured amplitudes, the distortion, the

noise, the phase, or the crosstalk value.  It returns the pairs of the

bin number and the selected data type listed below.

Data Types

(1) Level:

The function returns the measured amplitudes of the last received

multitone bins.

(2) Distortion:

The function returns the distortion values of all the bands between 20Hz

and 20kHz.  The distortion vector always starts with the bin number

greater than or equal to 20Hz, followed by the RSS (Root Sum Square)

value of the distortion from this bin to the first set signal bin.

Second pair is the bin number of the first signal bin and the RSS value

of the distortion bin between signal bin 1 and signal bin 2.  Last pair

is the last set signal bin with the RSS distortions from the last signal

bin to the last bin less than or equal to 20kHz.

The last returned x value, representing the frequency, is always the

first bin above 20kHz.  If a band between two set signal bins is too

narrow to measure distortions, then value "NaN" is returned indicating

that it is not a number.

(3) Noise:

The function returns the noise values of all the bands between 20Hz and

20kHz.  The noise vector always starts with the bin number greater than

or equal to 20Hz, followed by the RSS (Root Sum Square) value of the

noise from this bin to the first set signal bin.  Second pair is the bin

number of the first signal bin and the RSS value of the noise bin between

signal bin 1 and signal bin 2.  Last pair is the last set signal bin with

the RSS noise from the last signal bin to the last bin less than or equal

to 20kHz.

The last returned x value, representing the frequency, is always the

first bin above 20kHz.  If a band between two set signal bins is too

narrow to measure distortions, then value "NaN" is returned indicating

that it is not a number.

(4) Phase:

The function returns the measured phase-difference of the last measured

multitone signal.  The phase value can only be calculated if both

channels have recognized a trigger and if at least one bin of the

multitone signal is identical, i.e. is set at the same frequency on both

channels.  All answers are given in a string variable.  In order to get

accurate phase values, it is important to set the input ranges of both

channels to the same level.

(5) Crosstalk:

The function returns the measured crosstalk of the last received

multitone signal.  The crosstalk-value can only be calculated if at least

one bin of each channel is set at an exclusive frequency.  The channel

number defines the channel of the crosstalk measurement, e.g. if the

channel 1 is selected, the function measures all bins on channel 1 which

are only set on channel 2 and crosstalk at the same frequencies in

channel 1.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

selectChannel

Variable Type       ViInt16

This control selects the channel to be read.

Valid Range:

1 - Channel 1   (Default)

2 - Channel 2

dataType

Variable Type       ViInt16

This control selects the data type to be read for the selected

channel.

Valid Range:

0 - Level        (Default)

1 - Distortion

2 - Noise

3 - Phase

4 - Crosstalk

numberofBinsQueried

Variable Type       ViInt16 (passed by reference)

This output returns the total number of bins read from the selected

channel.

binNumber

Variable Type       ViInt16[]

This output returns an array of bin numbers.

Notes:

The maximum number of bins is 31.

amplitude

Variable Type       ViReal64[]

This output returns an array of the selected data for the selected

channel.

Notes:

The maximum number of bins is 31.

unit

Variable Type       ViChar[]

This output returns the unit of the selected data type.

Notes:

The string should contain at least 5 characters.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Select Channel) out of range.

BFFC0003  Parameter 3 (Data Type) out of range.

BFFC09F2  No trigger was detected.

BFFC09F3  Overflow occured at input.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_reset

ViStatus ntrt1m_reset (ViSession instrumentHandle);

Purpose

This function resets the instrument to a known state and sends

initialization commands to the instrument that set any necessary

programmatic variables such as Headers Off, Short Command form, and Data

Transfer Binary to the state necessary for the operation of the

instrument driver.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_revisionQuery

ViStatus ntrt1m_revisionQuery (ViSession instrumentHandle,

ViChar instrumentDriverRevision[],

ViChar firmwareRevision[]);

Purpose

This function returns the revision numbers of the instrument driver and

instrument firmware, and tells the user with which  instrument firmware

this revision of the driver is compatible.

This instrument driver's Revision Number is "Rev 1.0, 12/97, CVI 4.0.1"

Instrument Firmware Revision is unavailable.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

instrumentDriverRevision

Variable Type       ViChar[]

This control returns the Instrument Driver Software Revision.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

firmwareRevision

Variable Type       ViChar[]

This control returns the Instrument Firmware Revision.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_selfTest

ViStatus ntrt1m_selfTest (ViSession instrumentHandle, ViInt16 *selfTestResult,

ViChar selfTestMessage[]);

Purpose

This function runs the instrument's self test routine and returns the

test result(s).

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

selfTestResult

Variable Type       ViInt16 (passed by reference)

This control contains the value returned from the instrument self

test.  Zero means success.  For any other code, see the device's

operator's manual.

selfTestMessage

Variable Type       ViChar[]

This control contains the string returned from the self test. See the

device's operation manual for an explanation of the string's

contents.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_simpleMeas

ViStatus ntrt1m_simpleMeas (ViSession instrumentHandle, ViInt16 selectChannel,

ViReal64 inputRange, ViReal64 outputLevel,

ViBoolean preferredUnit, ViInt16 activeSignal,

ViInt16 dataType, ViInt16 *numberofBinsQueried,

ViInt16 binNumber[], ViReal64 amplitude[],

ViChar unit[]);

Purpose

This function sets the input range and the output level with the

preferred unit, selects the active signal, and reads either the measured

amplitudes, the distortion, the noise, the phase, or the crosstalk value.

It returns the pairs of the bin number and the selected data type listed

below.

Data Types

(1) Level:

The function returns the measured amplitudes of the last received

multitone bins.

(2) Distortion:

The function returns the distortion values of all the bands between 20Hz

and 20kHz.  The distortion vector always starts with the bin number

greater than or equal to 20Hz, followed by the RSS (Root Sum Square)

value of the distortion from this bin to the first set signal bin.

Second pair is the bin number of the first signal bin and the RSS value

of the distortion bin between signal bin 1 and signal bin 2.  Last pair

is the last set signal bin with the RSS distortions from the last signal

bin to the last bin less than or equal to 20kHz.

The last returned x value, representing the frequency, is always the

first bin above 20kHz.  If a band between two set signal bins is too

narrow to measure distortions, then value "NaN" is returned indicating

that it is not a number.

(3) Noise:

The function returns the noise values of all the bands between 20Hz and

20kHz.  The noise vector always starts with the bin number greater than

or equal to 20Hz, followed by the RSS (Root Sum Square) value of the

noise from this bin to the first set signal bin.  Second pair is the bin

number of the first signal bin and the RSS value of the noise bin between

signal bin 1 and signal bin 2.  Last pair is the last set signal bin with

the RSS noise from the last signal bin to the last bin less than or equal

to 20kHz.

The last returned x value, representing the frequency, is always the

first bin above 20kHz.  If a band between two set signal bins is too

narrow to measure distortions, then value "NaN" is returned indicating

that it is not a number.

(4) Phase:

The function returns the measured phase-difference of the last measured

multitone signal.  The phase value can only be calculated if both

channels have recognized a trigger and if at least one bin of the

multitone signal is identical, i.e. is set at the same frequency on both

channels.  All answers are given in a string variable.  In order to get

accurate phase values, it is important to set the input ranges of both

channels to the same level.

(5) Crosstalk:

The function returns the measured crosstalk of the last received

multitone signal.  The crosstalk-value can only be calculated if at least

one bin of each channel is set at an exclusive frequency.  The channel

number defines the channel of the crosstalk measurement, e.g. if the

channel 1 is selected, the function measures all bins on channel 1 which

are only set on channel 2 and crosstalk at the same frequencies in

channel 1.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

selectChannel

Variable Type       ViInt16

This control select the input channel to be configured.

Valid Range:

1 - Channel 1   (Default)

2 - Channel 2

inputRange

Variable Type       ViReal64

This control sets input range of the selected channel(s).

Valid Range:   Dependent upon the selected unit:

-60   to +20 dBVp

0.001 to  10 Vp

Default value:  0 dBVp

outputLevel

Variable Type       ViReal64

This control sets the output level of the selected channel(s).

Valid Range:   Dependent upon the selected unit:

-60   to +20 dBVp

0.001 to  10 Vp

Default value:  0 dBVp

preferredUnit

Variable Type       ViBoolean

This control selects the input unit for the selected channel(s).

Valid Range:

VI_OFF (0) - dBVp (Default)

VI_ON  (1) - Vp

activeSignal

Variable Type       ViInt16

This control selects the active multitone signal.

Valid Range:

1 - Signal in Memory 1

2 - Signal in Memory 2

3 - Signal in Memory 3

4 - Signal in Memory 4

Default Value:  1

dataType

Variable Type       ViInt16

This control selects the data type to be read for the selected

channel.

Valid Range:

0 - Level        (Default)

1 - Distortion

2 - Noise

3 - Phase

4 - Crosstalk

numberofBinsQueried

Variable Type       ViInt16 (passed by reference)

This output returns the total number of bins read from the selected

channel.

binNumber

Variable Type       ViInt16[]

This output returns an array of bin numbers.

Notes:

The maximum number of bins is 31.

amplitude

Variable Type       ViReal64[]

This output returns an array of the selected data for the selected

channel.

Notes:

The maximum number of bins is 31.

unit

Variable Type       ViChar[]

This output returns the unit of the selected data type.

Notes:

The string should contain at least 5 characters.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Select Channel) out of range.

BFFC0003  Parameter 3 (Input Range) out of range.

BFFC0004  Parameter 4 (Output Level)out of range.

BFFC0005  Parameter 5 (Preferred Unit) out of range.

BFFC0006  Parameter 6 (Active Signal) out of range.

BFFC0007  Parameter 7 (Data Type)out of range.

BFFC09F2  No trigger was detected.

BFFC09F3  Overflow occured at input.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_startActiveMtone

ViStatus ntrt1m_startActiveMtone (ViSession instrumentHandle);

Purpose

This function causes the generator to start with the active multitone

signal.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

ntrt1m_writeInstrData

ViStatus ntrt1m_writeInstrData (ViSession instrumentHandle,

ViString writeBuffer);

Purpose

This function writes commands and queries to the instrument to modify

parameters and query device settings.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

writeBuffer

Variable Type       ViString

The user can use this control to send common commands and queries to

the instrument. This control can also be used to write any valid

command to the instrument.

Default Value:  "*RST"

Notes:

(1) The command or query to be sent to the instrument may be a

literal enclosed in double quotes i.e. "*RST" or may be contained in

a variable of type string.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.